## Macro
FROM: | Mark Biegert
SUBJECT: | G23 Macro RIC Minimum Required SOC
DATE: | 22-Aug-2022
Manufacturing/Shipment/Fitting Change
Manufacturing Times
| Facility​ | Steps ​ | HA power on​ (minutes) | Profile | Current (mA) | Extended Charge (mAh) | SoC min requirement (%)​ | Comment
| Mfg Process​ | Assembly (w/ receiver)​ | 10 | Normal | 0.535 | 0.089167 | ​
|  | Brand​ | 5 | Normal | 0.535 | 0.044583 | ​
|  | WMTS​ | 10 | Normal | 0.535 | 0.089167 | 20​
|  | Acoustic (w/ receiver)​ | 10 | Normal | 0.535 | 0.6 | 20​ | Saif indicated acoustic testing took 2% SOC.
|  | Final Program​ | 5 | FFU | 0.865 | 0.072083 | ​
|  | FW upgrade ​ | 5 | FFU | 0.865 | 0.072083 | ​
| GDC​ | Brand & FP​ | 5 | FFU | 0.865 | 0.072083 | ​
| Time Total |  | 50 |  |  | 0.89975
Load Current Baseline
| Parameter Name | Symbol | Value | Unit | Description | Provenance
| Nominal Battery Current | _Inom | 0.535 | mA | Nominal Battery Current | DeviceCurrents workbook Sys Test
| D Firmware Update | _IFFU | 0.33 | mA | Increment for programming | DeviceCurrents workbook Sys Test
| HPM10 leakage | _I_HPM10 | 150 | nA | Leakage in the Power Down State | Mean HPM10 DSEN
| Minimum Final Config Volt | _V_fconfig | 1.205 | V | Standard Voltage Limit at Final Config | Adrian P.
Calculation Chain
| Budget Item | Symbol | Value | Unit | SOC State (Start) | Description | Provenance
| Battery Charge Rating | _QV | 30 | mAh |  | Capacity of the Battery | Varta Specification
| Starkey Range Correction | _DQ | 0.5 | mAh |  | Starkey's Operating range is different than the vendor specification because of limitations with the wireless hardware.
| Effective Mean Capacity | _Q | 29.5 | mAh |  | Bat Capacity over HA Voltage Range | Link
| Manufacturing Charge | _Qmfg | 0.89975 | mAh |  | Estimate of mfg charge | See Manufacturing Time Table
| Manufacturing SOC % | _SOCman | 3.05% |  | 49.38% | Starting SOC | See Mfg time table above
| Shelf Time | _Tshelf | 120 | days |  | Shipment and shelf time | MASTER-42226
| Manufacturing Time | _Tman | 14 | days |  | Wait time in Mfg | Adrian P.
| Shelf Current | _Ihpm10 | 150 | nA |  | Worst-case HPM10 | Mean HPM10 DSEN
| Self Discharge Rate | _kSDR | 10% | per year |  | Battery specification | Varta Specification
| HA Leakage | _Eleak | 0.4824 | mAh |  |  | Calculation
| SDR_Loss | _Esdr | 1.083 | mAh |  |  | Calculation
| Total Storage Losses | _Eshelf | 1.5654 | mAh |  |  | Calculation
| Storage SOC | _SOCstorage | 5.306% |  | 46.33% | SOC lost during storage | Calculation with Mfg Time
| Required SOC | _SOCmin | 33.03% |  | 33.03% | SOC @ Min Configuration Voltage | Reflects min config voltage
| Final Config Charge | _Qfinal | 0.21625 | mAh |  | Charge lost during final config | Calculation
| Final Configuration SOC | _Finalconfig | 0.7331% |  | 33.03% | Configure for final shipment | Assumes time in Mfg table
|  | _Tshelf2 | 120 | days
|  | _Eleak2 | 0.432 | mAh
|  | _Esdr2 | 0.96986 | mAh
|  | _Eshelf2 | 1.4019 | mAh
| Second Storage SOC | _SOCstorage2 | 4.752% |  | 32.29% | SOC lost during storage phase 2 | No Mfg time in this storage
| Fitting Time | _Tfitting | 3000% | min |  |  | MASTER-42226
| Fitting Current | _Ifit | 86.5% | mA |  | Assume we are always programming
| Fitting Charge | _Qfit | 43.25% | mAh
| Fitting SOC | _SOCfit | 1.466% |  | 27.54% | SOC lost during fitting | Calculation
| Battery Variation | _SOCerror | 8% |  | 26.08% | All bat calcs are mean. We need to include error bound. | Measured variation
| Min SOC at Mfg Start |  | 49.38% |  |  | Sum of losses | Computation
| Minimum battery voltage |  | 3.6834 | V |  | Battery voltage at mfg start | Computation
Low Battery Warning Threshold
| FFU SOC | _SOClowbat | 10% |  | Planned low battery warning SOC. Programming is precluded if low battery warning is on. | MASTER-75164
| Minimum Test SOC |  | 20% |  | Amount of charge at the start of test | Internal Discussion
System Constraints
| Budget Item | Symbol | Value | Unit | Description | Provenance
| Required SOC at Config | _SOCmin | 33.03% |  | SOC @ Minimum config voltage | Reflects min config voltage of 1.205 V
| FFU SOC | _SOClowbat | 10% |  | Planned low bat warning SOC. Programming is precluded if low battery warning is on. | MASTER-75164
Mfg and Config
14 days
120 days
Distribution Center
Final Config
120 days
Shipping
Fitting
1.205 V?
Fail
Note on F33: Mark Biegert: Worst-case HPM10 Value
Note on F43: Mark Biegert: The minimum SOC we can start at in production and have sufficient charge to meet all requirements by the time the customer gets the hearing aid.
Note on F52: Mark Biegert: This is set by the 1.205 V threshold requirement. This actually drives the production SOC requirement.
Note on G53: Mark Biegert: Sum of Final program, FW upgrade, and Brand and Final Program
Note on F54: Mark Biegert: Must equal, at a minimum, the required config voltage
Note on F63: Mark Biegert: The SOC expected to be in that battery at the completion of the fitting session.
Note on F77: Mark Biegert: I am using 10% instead of an actual voltage value because we are planning on moving to a SOC-based low-voltage threshold in the future (following the model of cell phones). The current voltage-based threshold corresponds to ~8% SOC, so this 10% level provides a slightly higher required manufacturing SOC.

## Micro
FROM: | Mark Biegert
SUBJECT: | G23 Micro RIC Minimum Required SOC
DATE: | 22-Aug-2022
Manufacturing/Shipment/Fitting Change
Manufacturing Times
| Facility​ | Steps ​ | HA power on​ (minutes) | Profile | Current (mA) | Extended Charge (mAh) | SoC min requirement (%)​ | Comment
| Mfg Process​ | Assembly (w/ receiver)​ | 10 | Normal | 0.44 | 0.073333 | ​
|  | Brand​ | 5 | Normal | 0.44 | 0.036667 | ​
|  | WMTS​ | 10 | Normal | 0.44 | 0.073333 | 20​
|  | Acoustic (w/ receiver)​ | 10 | Normal | 0.44 | 0.4 | 20​ | Saif indicated acoustic testing took 2% SOC.
|  | Final Program​ | 5 | FFU | 0.77 | 0.064167 | ​
|  | FW upgrade ​ | 5 | FFU | 0.77 | 0.064167 | ​
| GDC​ | Brand & FP​ | 5 | FFU | 0.77 | 0.064167 | ​
| Time Total |  | 50 |  |  | 0.77583
Load Current Baseline
| Parameter Name | Symbol | Value | Unit | Description | Provenance
| Nominal Battery Current | _Inom | 0.44 | mA | Nominal Battery Current | DeviceCurrents workbook Sys Test
| D Firmware Update | _IFFU | 0.33 | mA | Increment for programming | DeviceCurrents workbook Sys Test
| HPM10 leakage | _I_HPM10 | 150 | nA | Leakage in the Power Down State | Mean HPM10 DSEN
| Minimum Final Config Volt | _V_fconfig | 1.205 | V | Standard Voltage Limit at Final | Adrian P.
Calculation Chain
| Budget Item | Symbol | Value | Unit | SOC State (Start) | Description | Provenance
| Battery Charge Rating | _QV | 20 | mAh |  | Capacity of the Battery | Varta Specification
| Starkey Range Correction | _DQ | 0.5 | mAh |  | Starkey's Operating range is different than the vendor specification because of limitations with the wireless hardware. | Battery Parameter Detail
| Effective Mean Capacity | _Q | 19.5 | mAh |  | Bat Capacity over HA Voltage Range | Link
| Manufacturing Charge | _Qmfg | 0.77583 | mAh |  | Estimate of mfg charge | See Manufacturing Time Table
| Manufacturing SOC % | _SOCman | 3.979% |  | 51.15% | Starting SOC | SOC used in Mfg
| Shelf Time | _Tshelf | 120 | days |  | Shipment and shelf time | MASTER-42226
| Manufacturing Time | _Tman | 14 | days |  | Wait time in Mfg | Adrian P.
| Shelf Current | _Ihpm10 | 150 | nA |  | Worst-case HPM10 | Mean HPM10 DSEN
| Self Discharge Rate | _kSDR | 10% | per year |  | Battery specification | Varta Specification
| HA Leakage | _Eleak | 0.4824 | mAh |  |  | Calculation
| SDR_Loss | _Esdr | 0.71589 | mAh |  |  | Calculation
| Total Storage Losses | _Eshelf | 1.1983 | mAh |  |  | Calculation
| Storage SOC | _SOCstorage | 6.145% |  | 47.17% | SOC lost during storage | Calculation with Mfg Time
| Required SOC | _SOCmin | 33.03% |  | 33.03% | SOC @ Min Configuration Voltage | Reflects min config voltage
| Final Config Charge | _Qfinal | 0.1925 | mAh
| Final Configuration SOC | _Finalconfig | 0.9872% |  | 33.03%
|  | _Tshelf2 | 120 | days
|  | _Eleak2 | 0.432 | mAh
|  | _Esdr2 | 0.6411 | mAh
|  | _Eshelf2 | 1.0731 | mAh
| Second Storage SOC | _SOCstorage2 | 5.503% |  | 32.04% | SOC lost during storage phase 2 | No Mfg time in this storage
| Fitting Time | _Tfitting | 30 | min |  |  | MASTER-42226
| Fitting Current | _Ifit | 0.77 | mA |  | Assume we are always programming
| Fitting Charge | _Qfit | 0.385 | mAh
| Fitting SOC | _SOCfit | 1.974% |  | 26.54% | SOC lost during fitting | Calculation
| Battery Variation | _SOCerror | 8% |  | 24.56% | All bat calcs are mean. We need to include error bound. | Measured variation
| Min SOC at Mfg Start |  | 51.15% |  |  | Sum of losses | Computation
| Minimum battery voltage |  | 3.6955 | V |  | Battery voltage at mfg start | Computation
Low Battery Warning Threshold
| FFU SOC | _SOClowbat | 10% |  | Planned low battery warning SOC. Programming is precluded if low battery warning is on. | MASTER-75164
| Minimum Test SOC |  | 20% |  | Amount of charge at the start of test | Internal Discussion
System Constraints
| Budget Item | Symbol | Value | Unit | Description | Provenance
| Required SOC at Config | _SOCmin | 33.03% |  | SOC @ Minimum config voltage | Reflects min config voltage
| FFU SOC | _SOClowbat | 10% |  | Planned low bat warning SOC. Programming is precluded if low battery warning is on. | MASTER-75164
Note on F33: Mark Biegert: Worst-case HPM10 value
Note on F43: Mark Biegert: The minimum SOC that we can charge the battery to and meet all requirements.
Note on F63: Mark Biegert: SOC expected to be in the battery on the completion of the fitting session.
Note on G77: Mark Biegert: I am using 10% instead of an actual voltage value because we are planning on moving to a SOC-based low-voltage threshold in the future (following the model of cell phones). The current voltage-based threshold corresponds to ~8% SOC, so this 10% level provides a slightly higher required manufacturing SOC.
Mfg and Config
14 days
120 days
Distribution Center
Final Config
120 days
Shipping
Fitting
1.205 V?
Fail

## Varta_SOC
FROM: | Mark Biegert
SUBJECT: | Varta SOC Curve
DATE: | 22-Aug-2022
Provenance
Polarion Battery SOC Curve Document
Varta SOC
| SOC | VBAT | VHA
| 0% | 3.414 | 1.138
| 1% | 3.4277 | 1.1426
| 2% | 3.4371 | 1.1457
| 3% | 3.4441 | 1.148
| 4% | 3.4501 | 1.15
| 5% | 3.4548 | 1.1516
| 6% | 3.4596 | 1.1532
| 7% | 3.4647 | 1.1549
| 8% | 3.4698 | 1.1566
| 9% | 3.4779 | 1.1593
| 10% | 3.4871 | 1.1624
| 11% | 3.4977 | 1.1659
| 12% | 3.5074 | 1.1691
| 13% | 3.5155 | 1.1718
| 14% | 3.5239 | 1.1746
| 15% | 3.532 | 1.1773
| 16% | 3.538 | 1.1793
| 17% | 3.5442 | 1.1814
| 18% | 3.5504 | 1.1835 | Chart Title | Varta Accu60/94 VHA versus SOC | Threshold |  |  | VA | [5 empty] | OZ | [5 empty] | VZ |  |  | ZA
| 19% | 3.5559 | 1.1853 | X Axis | SOC (%) |  | Threshold |  |  | Varta App SOC | [5 empty] | Original SOC | [5 empty] | SOC | VBAT |  | Z22A App SOC
| 20% | 3.5615 | 1.1872 | Y Axis | VHA (V) |  | 1.205 | 0 |  | 0,10,30,60,80,100 | [5 empty] | 1.16015625,1.19140625,1.21875,1.25390625,1.296875,1.31640625 | [5 empty] | 0% | 3 |  | 1.138,1.225,1.245,1.267,1.284,1.316
| 21% | 3.5678 | 1.1893 |  |  |  | 1.205 | 1 |  | 1.138,1.156,1.193,1.224,1.271,1.321 | [6 empty] | VHA | VBAT | SOC | SOC% |  | 1% | 3.27 |  |  | VHA | VBAT | SOC | SOC%
| 22% | 3.573 | 1.191 | [7 empty] | VHA | VBAT | SOC | SOC% |  |  | 1.1602 | 3.4805 | 0 | 0% |  | 2% | 3.39 |  |  | 1.138 | 3.414 | 0 | 0%
| 23% | 3.5776 | 1.1925 | [7 empty] | 1.138 | 3.414 | 0 | 0% |  |  | 1.1914 | 3.5742 | 10 | 10% |  | 3% | 3.455 |  |  | 1.225 | 3.675 | 10 | 10%
| 24% | 3.5821 | 1.194 | [7 empty] | 1.156 | 3.468 | 10 | 10% |  |  | 1.2188 | 3.6562 | 30 | 30% |  | 5% | 3.558 |  |  | 1.245 | 3.735 | 30 | 30%
| 25% | 3.5864 | 1.1955 |  |  |  | 3 |  |  |  | 1.193 | 3.579 | 30 | 30% |  |  | 1.2539 | 3.7617 | 60 | 60% |  | 10% | 3.658 |  |  | 1.267 | 3.801 | 60 | 60%
| 26% | 3.5906 | 1.1969 | [7 empty] | 1.224 | 3.672 | 60 | 60% |  |  | 1.2969 | 3.8906 | 80 | 80% |  | 15% | 3.676 |  |  | 1.284 | 3.852 | 80 | 80%
| 27% | 3.5942 | 1.1981 | [7 empty] | 1.271 | 3.813 | 80 | 80% |  |  | 1.3164 | 3.9492 | 100 | 100% |  | 20% | 3.708 |  |  | 1.316 | 3.948 | 100 | 100%
| 28% | 3.5978 | 1.1993 | [7 empty] | 1.321 | 3.963 | 100 | 100% | [7 empty] | 25% | 3.728 |  |  |  | 4 | 100 | 100%
| 29% | 3.6013 | 1.2004 | [8 empty] | 4.2 |  | 100% | [7 empty] | 30% | 3.743 |  |  |  | 4.2 | 100 | 100%
| 30% | 3.6046 | 1.2015 | [18 empty] | 35% | 3.757
| 31% | 3.6079 | 1.2026 | [18 empty] | 40% | 3.767
| 32% | 3.6113 | 1.2038 | [18 empty] | 50% | 3.797
| 33% | 3.6149 | 1.205 | [18 empty] | 55% | 3.817
| 34% | 3.6185 | 1.2062 | [18 empty] | 60% | 3.845
| 35% | 3.6218 | 1.2073 | [18 empty] | 65% | 3.876
| 36% | 3.6249 | 1.2083 | [18 empty] | 70% | 3.91
| 37% | 3.628 | 1.2093 | [18 empty] | 75% | 3.95
| 38% | 3.6314 | 1.2105 | [18 empty] | 80% | 3.988
| 39% | 3.6354 | 1.2118 | [18 empty] | 85% | 4.032
| 40% | 3.6393 | 1.2131 | [18 empty] | 90% | 4.075
| 41% | 3.6432 | 1.2144 | [18 empty] | 95% | 4.12
| 42% | 3.6472 | 1.2157 | [18 empty] | 100% | 4.18
| 43% | 3.6512 | 1.2171
| 44% | 3.655 | 1.2183
| 45% | 3.6589 | 1.2196
| 46% | 3.6646 | 1.2215
| 47% | 3.6708 | 1.2236
| 48% | 3.6756 | 1.2252
| 49% | 3.6806 | 1.2269
| 50% | 3.6879 | 1.2293
| 51% | 3.6945 | 1.2315
| 52% | 3.7012 | 1.2337
| 53% | 3.71 | 1.2367
| 54% | 3.7182 | 1.2394
| 55% | 3.7302 | 1.2434
| 56% | 3.7382 | 1.2461
| 57% | 3.7513 | 1.2504
| 58% | 3.759 | 1.253
| 59% | 3.7717 | 1.2572
| 60% | 3.779 | 1.2597
| 61% | 3.7872 | 1.2624
| 62% | 3.7972 | 1.2657
| 63% | 3.8059 | 1.2686
| 64% | 3.8155 | 1.2718
| 65% | 3.8249 | 1.275
| 66% | 3.8338 | 1.2779
| 67% | 3.8433 | 1.2811
| 68% | 3.8529 | 1.2843
| 69% | 3.863 | 1.2877
| 70% | 3.8709 | 1.2903
| 71% | 3.8817 | 1.2939
| 72% | 3.8892 | 1.2964
| 73% | 3.9003 | 1.3001
| 74% | 3.9079 | 1.3026
| 75% | 3.9166 | 1.3055
| 76% | 3.9261 | 1.3087
| 77% | 3.9349 | 1.3116
| 78% | 3.9442 | 1.3147
| 79% | 3.9541 | 1.318
| 80% | 3.9631 | 1.321
| 81% | 3.9732 | 1.3244
| 82% | 3.9832 | 1.3277
| 83% | 3.9939 | 1.3313
| 84% | 4.0032 | 1.3344
| 85% | 4.0138 | 1.3379
| 86% | 4.0232 | 1.3411
| 87% | 4.034 | 1.3447
| 88% | 4.0439 | 1.348
| 89% | 4.0545 | 1.3515
| 90% | 4.0647 | 1.3549
| 91% | 4.0754 | 1.3585
| 92% | 4.0859 | 1.362
| 93% | 4.0977 | 1.3659
| 94% | 4.1082 | 1.3694
| 95% | 4.1213 | 1.3738
| 96% | 4.1321 | 1.3774
| 97% | 4.1444 | 1.3815
| 98% | 4.157 | 1.3857
| 99% | 4.1709 | 1.3903
| 100% | 4.2 | 1.4
Note: This is the Varta measured mean SOC curve. This is different from the SOC curve loaded into the FW. The SOC curve in FW differs at high battery voltages to deal with different battery relaxation voltages. I deal with the battery variation by assigning a variation "tax" to the mean value.

## Z22A_SOC
FROM: | Mark Biegert
SUBJECT: | Measured Mean Z22A SOC Curve
DATE: | 30-Nov-2022
| Time | Current | Voltage | SOC | VHA
| 37.7333 | 0.67001 | 3.413 | 0 | 1.1377
| 37.7167 | 0.67715 | 3.414 | 0 | 1.138
| 37.7167 | 0.67715 | 3.4167 | 0.00027417 | 1.1389
| 37.7 | 0.67715 | 3.417 | 0.0003373 | 1.139
| 37.7 | 0.67358 | 3.4207 | 0.00071594 | 1.1402
| 37.6833 | 0.67715 | 3.4209 | 0.00077923 | 1.1403
| 37.6833 | 0.67715 | 3.4247 | 0.0011577 | 1.1416
| 37.6667 | 0.67715 | 3.4248 | 0.0012212 | 1.1416
| 37.65 | 0.67715 | 3.4286 | 0.0016631 | 1.1429
| 37.6667 | 0.67715 | 3.4287 | 0.0015995 | 1.1429
| 37.6333 | 0.67358 | 3.4321 | 0.002105 | 1.144
| 37.65 | 0.67358 | 3.4324 | 0.0020412 | 1.1441
| 37.6167 | 0.67358 | 3.4358 | 0.002547 | 1.1453
| 37.6333 | 0.67358 | 3.436 | 0.002483 | 1.1453
| 37.6 | 0.67358 | 3.4395 | 0.0029889 | 1.1465
| 37.6167 | 0.67358 | 3.4396 | 0.0029248 | 1.1465
| 37.5833 | 0.67715 | 3.4432 | 0.0034309 | 1.1477
| 37.6 | 0.67358 | 3.4433 | 0.0033666 | 1.1478
| 37.5833 | 0.67715 | 3.4467 | 0.0038083 | 1.1489
| 37.5667 | 0.67358 | 3.4468 | 0.0038728 | 1.1489
| 37.55 | 0.67715 | 3.4503 | 0.0043147 | 1.1501
| 37.5667 | 0.67001 | 3.4505 | 0.0042501 | 1.1502
| 37.5333 | 0.67715 | 3.4537 | 0.0047567 | 1.1512
| 37.55 | 0.67358 | 3.4538 | 0.0046919 | 1.1513
| 37.5167 | 0.67715 | 3.4569 | 0.0051986 | 1.1523
| 37.5333 | 0.67358 | 3.4573 | 0.0051336 | 1.1524
| 37.5 | 0.67358 | 3.4601 | 0.0056405 | 1.1534
| 37.5167 | 0.67715 | 3.4607 | 0.0055754 | 1.1536
| 37.4833 | 0.67715 | 3.4633 | 0.0060825 | 1.1544
| 37.5 | 0.67715 | 3.4641 | 0.0060172 | 1.1547
| 37.4667 | 0.67358 | 3.4664 | 0.0065244 | 1.1555
| 37.4833 | 0.67715 | 3.4672 | 0.0064589 | 1.1557
| 37.45 | 0.67358 | 3.4697 | 0.0069664 | 1.1566
| 37.4667 | 0.67358 | 3.4706 | 0.0069007 | 1.1569
| 37.4333 | 0.67358 | 3.4729 | 0.0074083 | 1.1577
| 37.45 | 0.67358 | 3.4735 | 0.0073425 | 1.1578
| 37.4167 | 0.67358 | 3.4759 | 0.0078502 | 1.1586
| 37.4333 | 0.67715 | 3.4765 | 0.0077843 | 1.1588
| 37.4 | 0.67358 | 3.4791 | 0.0082922 | 1.1597
| 37.4167 | 0.67358 | 3.4796 | 0.008226 | 1.1599
| 37.3833 | 0.67358 | 3.4821 | 0.0087341 | 1.1607
| 37.4 | 0.67358 | 3.4826 | 0.0086678 | 1.1609
| 37.3667 | 0.67358 | 3.4849 | 0.009176 | 1.1616
| 37.3833 | 0.67715 | 3.4857 | 0.0091096 | 1.1619
| 37.35 | 0.67715 | 3.4877 | 0.009618 | 1.1626
| 37.3667 | 0.67358 | 3.4888 | 0.0095513 | 1.1629
| 37.3333 | 0.67715 | 3.4906 | 0.01006 | 1.1635
| 37.35 | 0.67358 | 3.4917 | 0.0099931 | 1.1639
| 37.3167 | 0.67001 | 3.4935 | 0.010502 | 1.1645
| 37.3333 | 0.67715 | 3.4945 | 0.010435 | 1.1648
| 37.3 | 0.67358 | 3.4962 | 0.010944 | 1.1654
| 37.3167 | 0.67358 | 3.4974 | 0.010877 | 1.1658
| 37.2833 | 0.67715 | 3.499 | 0.011386 | 1.1663
| 37.3 | 0.67358 | 3.5002 | 0.011318 | 1.1667
| 37.2667 | 0.67715 | 3.5017 | 0.011828 | 1.1672
| 37.2833 | 0.67715 | 3.5029 | 0.01176 | 1.1676
| 37.25 | 0.68072 | 3.5045 | 0.01227 | 1.1682
| 37.2667 | 0.67715 | 3.5055 | 0.012202 | 1.1685
| 37.2333 | 0.67358 | 3.5072 | 0.012712 | 1.1691
| 37.25 | 0.67358 | 3.5081 | 0.012644 | 1.1694
| 37.2167 | 0.67715 | 3.5098 | 0.013153 | 1.1699
| 37.2333 | 0.67715 | 3.5107 | 0.013085 | 1.1702
| 37.2 | 0.67001 | 3.5124 | 0.013595 | 1.1708
| 37.2167 | 0.67358 | 3.5134 | 0.013527 | 1.1711
| 37.1833 | 0.67715 | 3.5148 | 0.014037 | 1.1716
| 37.2 | 0.67358 | 3.5162 | 0.013969 | 1.1721
| 37.1667 | 0.67358 | 3.5171 | 0.014479 | 1.1724
| 37.1833 | 0.67358 | 3.5187 | 0.014411 | 1.1729
| 37.15 | 0.67358 | 3.5196 | 0.014921 | 1.1732
| 37.1667 | 0.67358 | 3.5213 | 0.014853 | 1.1738
| 37.1333 | 0.67715 | 3.5222 | 0.015363 | 1.1741
| 37.15 | 0.67715 | 3.5236 | 0.015294 | 1.1745
| 37.1167 | 0.67358 | 3.5247 | 0.015805 | 1.1749
| 37.1333 | 0.67715 | 3.526 | 0.015736 | 1.1753
| 37.1 | 0.67358 | 3.5272 | 0.016247 | 1.1757
| 37.1167 | 0.67358 | 3.5284 | 0.016178 | 1.1761
| 37.0833 | 0.67358 | 3.5296 | 0.016689 | 1.1765
| 37.1 | 0.67358 | 3.5307 | 0.01662 | 1.1769
| 37.0667 | 0.67358 | 3.5319 | 0.017131 | 1.1773
| 37.0833 | 0.67358 | 3.533 | 0.017061 | 1.1777
| 37.05 | 0.67715 | 3.5342 | 0.017573 | 1.1781
| 37.0667 | 0.67358 | 3.5353 | 0.017503 | 1.1784
| 37.0333 | 0.67001 | 3.5364 | 0.018015 | 1.1788
| 37.05 | 0.67715 | 3.5377 | 0.017945 | 1.1792
| 37.0167 | 0.67358 | 3.5386 | 0.018457 | 1.1795
| 37.0333 | 0.67358 | 3.5399 | 0.018387 | 1.18
| 37 | 0.67715 | 3.5408 | 0.018899 | 1.1803
| 37.0167 | 0.67358 | 3.5423 | 0.018829 | 1.1808
| 36.9833 | 0.67358 | 3.5431 | 0.019341 | 1.181
| 37 | 0.67358 | 3.5446 | 0.01927 | 1.1815
| 36.9667 | 0.67715 | 3.5454 | 0.019783 | 1.1818
| 36.9833 | 0.67715 | 3.547 | 0.019712 | 1.1823
| 36.95 | 0.67358 | 3.5477 | 0.020224 | 1.1826
| 36.9667 | 0.67358 | 3.5488 | 0.020154 | 1.1829
| 36.9333 | 0.67715 | 3.5498 | 0.020666 | 1.1833
| 36.95 | 0.67715 | 3.5509 | 0.020596 | 1.1836
| 36.9167 | 0.67001 | 3.5522 | 0.021108 | 1.1841
| 36.9333 | 0.67358 | 3.5531 | 0.021037 | 1.1844
| 36.9 | 0.67358 | 3.5541 | 0.02155 | 1.1847
| 36.9167 | 0.67715 | 3.555 | 0.021479 | 1.185
| 36.8833 | 0.67715 | 3.556 | 0.021992 | 1.1853
| 36.9 | 0.67715 | 3.557 | 0.021921 | 1.1857
| 36.8667 | 0.67715 | 3.558 | 0.022434 | 1.186
| 36.8833 | 0.67715 | 3.559 | 0.022363 | 1.1863
| 36.85 | 0.67358 | 3.5602 | 0.022876 | 1.1867
| 36.8667 | 0.67715 | 3.5611 | 0.022804 | 1.187
| 36.8333 | 0.67001 | 3.5623 | 0.023318 | 1.1874
| 36.85 | 0.67715 | 3.5631 | 0.023246 | 1.1877
| 36.8167 | 0.67715 | 3.5644 | 0.02376 | 1.1881
| 36.8333 | 0.67358 | 3.5653 | 0.023688 | 1.1884
| 36.8 | 0.67358 | 3.5664 | 0.024202 | 1.1888
| 36.8167 | 0.67358 | 3.5673 | 0.02413 | 1.1891
| 36.7833 | 0.67358 | 3.5684 | 0.024644 | 1.1895
| 36.8 | 0.67001 | 3.5692 | 0.024572 | 1.1897
| 36.7667 | 0.67358 | 3.5704 | 0.025086 | 1.1901
| 36.7833 | 0.67715 | 3.5711 | 0.025013 | 1.1904
| 36.75 | 0.67715 | 3.5723 | 0.025528 | 1.1908
| 36.7667 | 0.67715 | 3.5729 | 0.025455 | 1.191
| 36.7333 | 0.67358 | 3.574 | 0.02597 | 1.1913
| 36.75 | 0.67358 | 3.5749 | 0.025897 | 1.1916
| 36.7167 | 0.67715 | 3.576 | 0.026412 | 1.192
| 36.7333 | 0.67715 | 3.5766 | 0.026339 | 1.1922
| 36.7 | 0.67358 | 3.5778 | 0.026854 | 1.1926
| 36.7167 | 0.67358 | 3.5783 | 0.02678 | 1.1928
| 36.6833 | 0.67358 | 3.5796 | 0.027295 | 1.1932
| 36.7 | 0.67358 | 3.5802 | 0.027222 | 1.1934
| 36.6667 | 0.67358 | 3.5814 | 0.027737 | 1.1938
| 36.6833 | 0.67715 | 3.5821 | 0.027664 | 1.194
| 36.65 | 0.67358 | 3.5833 | 0.028179 | 1.1945
| 36.6667 | 0.67715 | 3.584 | 0.028106 | 1.1947
| 36.6333 | 0.67358 | 3.5853 | 0.028621 | 1.1951
| 36.65 | 0.67358 | 3.5858 | 0.028547 | 1.1953
| 36.6167 | 0.67358 | 3.587 | 0.029063 | 1.1957
| 36.6333 | 0.67715 | 3.5876 | 0.028989 | 1.1959
| 36.6 | 0.67358 | 3.5888 | 0.029505 | 1.1963
| 36.6167 | 0.67358 | 3.5895 | 0.029431 | 1.1965
| 36.5833 | 0.67358 | 3.5904 | 0.029947 | 1.1968
| 36.6 | 0.67715 | 3.591 | 0.029873 | 1.197
| 36.5667 | 0.67358 | 3.5921 | 0.030389 | 1.1974
| 36.5833 | 0.67358 | 3.5927 | 0.030315 | 1.1976
| 36.55 | 0.67715 | 3.5939 | 0.030831 | 1.198
| 36.5667 | 0.67358 | 3.5943 | 0.030756 | 1.1981
| 36.5333 | 0.67358 | 3.5954 | 0.031273 | 1.1985
| 36.55 | 0.67358 | 3.5961 | 0.031198 | 1.1987
| 36.5167 | 0.67715 | 3.5972 | 0.031715 | 1.1991
| 36.5333 | 0.67001 | 3.5977 | 0.03164 | 1.1992
| 36.5 | 0.67358 | 3.5989 | 0.032157 | 1.1996
| 36.5167 | 0.67715 | 3.5997 | 0.032082 | 1.1999
| 36.4833 | 0.67715 | 3.6008 | 0.032599 | 1.2003
| 36.5 | 0.67715 | 3.6014 | 0.032523 | 1.2005
| 36.4667 | 0.67358 | 3.6023 | 0.033041 | 1.2008
| 36.4833 | 0.67358 | 3.603 | 0.032965 | 1.201
| 36.45 | 0.67358 | 3.604 | 0.033483 | 1.2013
| 36.4667 | 0.67358 | 3.6045 | 0.033407 | 1.2015
| 36.4333 | 0.67715 | 3.6054 | 0.033925 | 1.2018
| 36.45 | 0.67715 | 3.6061 | 0.033849 | 1.202
| 36.4167 | 0.67358 | 3.607 | 0.034366 | 1.2023
| 36.4333 | 0.67715 | 3.6076 | 0.03429 | 1.2025
| 36.4 | 0.67715 | 3.6084 | 0.034808 | 1.2028
| 36.4167 | 0.67358 | 3.6091 | 0.034732 | 1.203
| 36.3833 | 0.67715 | 3.61 | 0.03525 | 1.2033
| 36.4 | 0.67358 | 3.6107 | 0.035174 | 1.2036
| 36.3667 | 0.67715 | 3.6117 | 0.035692 | 1.2039
| 36.3833 | 0.67358 | 3.6123 | 0.035616 | 1.2041
| 36.35 | 0.67358 | 3.6133 | 0.036134 | 1.2044
| 36.3667 | 0.67715 | 3.614 | 0.036058 | 1.2047
| 36.3333 | 0.67358 | 3.6149 | 0.036576 | 1.205
| 36.35 | 0.67358 | 3.6157 | 0.036499 | 1.2052
| 36.3167 | 0.67715 | 3.6165 | 0.037018 | 1.2055
| 36.3333 | 0.67358 | 3.6172 | 0.036941 | 1.2057
| 36.3 | 0.67358 | 3.618 | 0.03746 | 1.206
| 36.3167 | 0.67358 | 3.6187 | 0.037383 | 1.2063
| 36.2833 | 0.67358 | 3.6194 | 0.037902 | 1.2065
| 36.3 | 0.67358 | 3.6201 | 0.037825 | 1.2067
| 36.2667 | 0.67358 | 3.6205 | 0.038344 | 1.2068
| 36.2833 | 0.67715 | 3.6214 | 0.038266 | 1.2071
| 36.25 | 0.67358 | 3.622 | 0.038786 | 1.2073
| 36.2667 | 0.67715 | 3.6229 | 0.038708 | 1.2076
| 36.2333 | 0.67715 | 3.6233 | 0.039228 | 1.2078
| 36.25 | 0.67358 | 3.6242 | 0.03915 | 1.2081
| 36.2167 | 0.67715 | 3.6246 | 0.03967 | 1.2082
| 36.2333 | 0.67715 | 3.6255 | 0.039592 | 1.2085
| 36.2 | 0.67358 | 3.6258 | 0.040112 | 1.2086
| 36.2167 | 0.67358 | 3.6271 | 0.040033 | 1.209
| 36.1833 | 0.67715 | 3.6273 | 0.040554 | 1.2091
| 36.2 | 0.67358 | 3.6285 | 0.040475 | 1.2095
| 36.1667 | 0.67715 | 3.6286 | 0.040996 | 1.2095
| 36.1833 | 0.67715 | 3.63 | 0.040917 | 1.21
| 36.15 | 0.67358 | 3.6301 | 0.041437 | 1.21
| 36.1667 | 0.67358 | 3.6312 | 0.041359 | 1.2104
| 36.1333 | 0.67358 | 3.6316 | 0.041879 | 1.2105
| 36.15 | 0.67358 | 3.6327 | 0.041801 | 1.2109
| 36.1167 | 0.67715 | 3.6329 | 0.042321 | 1.211
| 36.1333 | 0.67358 | 3.6338 | 0.042242 | 1.2113
| 36.1 | 0.67715 | 3.634 | 0.042763 | 1.2113
| 36.1167 | 0.67358 | 3.6351 | 0.042684 | 1.2117
| 36.0833 | 0.67001 | 3.6352 | 0.043205 | 1.2117
| 36.1 | 0.67715 | 3.6363 | 0.043126 | 1.2121
| 36.0667 | 0.67715 | 3.6364 | 0.043647 | 1.2122
| 36.0833 | 0.67358 | 3.6375 | 0.043568 | 1.2125
| 36.0667 | 0.67001 | 3.6389 | 0.044009 | 1.213
| 36.0167 | 0.67715 | 3.6399 | 0.044973 | 1.2133
| 36.05 | 0.67715 | 3.6401 | 0.044451 | 1.2134
| 36 | 0.67715 | 3.6413 | 0.045415 | 1.2138
| 36.0333 | 0.67715 | 3.6414 | 0.044893 | 1.2138
| 35.9833 | 0.67358 | 3.6424 | 0.045857 | 1.2141
| 36.0167 | 0.67715 | 3.6426 | 0.045335 | 1.2142
| 36 | 0.67715 | 3.6437 | 0.045776 | 1.2146
| 35.95 | 0.67715 | 3.6448 | 0.046741 | 1.2149
| 35.9833 | 0.67715 | 3.6449 | 0.046218 | 1.215
| 35.9667 | 0.67715 | 3.6458 | 0.04666 | 1.2153
| 35.9167 | 0.67358 | 3.6469 | 0.047625 | 1.2156
| 35.95 | 0.67715 | 3.647 | 0.047102 | 1.2157
| 35.9 | 0.67715 | 3.648 | 0.048067 | 1.216
| 35.9333 | 0.67001 | 3.6481 | 0.047544 | 1.216
| 35.8833 | 0.67715 | 3.6488 | 0.048508 | 1.2163
| 35.9167 | 0.67358 | 3.6492 | 0.047985 | 1.2164
| 35.8667 | 0.67358 | 3.65 | 0.04895 | 1.2167
| 35.9 | 0.67358 | 3.6503 | 0.048427 | 1.2168
| 35.85 | 0.67715 | 3.6509 | 0.049392 | 1.217
| 35.8833 | 0.67715 | 3.6513 | 0.048869 | 1.2171
| 35.8333 | 0.67358 | 3.652 | 0.049834 | 1.2173
| 35.8667 | 0.67358 | 3.6524 | 0.049311 | 1.2175
| 35.8167 | 0.67715 | 3.6531 | 0.050276 | 1.2177
| 35.85 | 0.67715 | 3.6535 | 0.049752 | 1.2178
| 35.8 | 0.67715 | 3.6541 | 0.050718 | 1.218
| 35.8333 | 0.67358 | 3.6543 | 0.050194 | 1.2181
| 35.7833 | 0.67715 | 3.6551 | 0.05116 | 1.2184
| 35.8167 | 0.67715 | 3.6552 | 0.050636 | 1.2184
| 35.8 | 0.67715 | 3.6561 | 0.051078 | 1.2187
| 35.75 | 0.67358 | 3.6568 | 0.052044 | 1.2189
| 35.7833 | 0.67358 | 3.6569 | 0.051519 | 1.219
| 35.7333 | 0.67715 | 3.6577 | 0.052486 | 1.2192
| 35.7667 | 0.67358 | 3.6578 | 0.051961 | 1.2193
| 35.7167 | 0.67715 | 3.6586 | 0.052928 | 1.2195
| 35.75 | 0.67715 | 3.6587 | 0.052403 | 1.2196
| 35.7 | 0.67358 | 3.6597 | 0.05337 | 1.2199
| 35.7333 | 0.67715 | 3.6598 | 0.052845 | 1.2199
| 35.6833 | 0.67715 | 3.6603 | 0.053812 | 1.2201
| 35.7167 | 0.67358 | 3.6607 | 0.053287 | 1.2202
| 35.6667 | 0.67715 | 3.6613 | 0.054254 | 1.2204
| 35.7 | 0.67715 | 3.6616 | 0.053728 | 1.2205
| 35.65 | 0.67715 | 3.6622 | 0.054696 | 1.2207
| 35.6833 | 0.67358 | 3.6623 | 0.05417 | 1.2208
| 35.6667 | 0.67001 | 3.6629 | 0.054612 | 1.221
| 35.6333 | 0.67358 | 3.663 | 0.055138 | 1.221
| 35.65 | 0.67715 | 3.6638 | 0.055054 | 1.2213
| 35.6167 | 0.67715 | 3.6638 | 0.055579 | 1.2213
| 35.6333 | 0.67358 | 3.6644 | 0.055495 | 1.2215
| 35.6 | 0.67358 | 3.6644 | 0.056021 | 1.2215
| 35.6167 | 0.67358 | 3.6651 | 0.055937 | 1.2217
| 35.5667 | 0.67715 | 3.6658 | 0.056905 | 1.2219
| 35.6 | 0.67715 | 3.6659 | 0.056379 | 1.222
| 35.55 | 0.67715 | 3.6666 | 0.057347 | 1.2222
| 35.5833 | 0.67358 | 3.6667 | 0.056821 | 1.2222
| 35.5667 | 0.67715 | 3.6673 | 0.057262 | 1.2225
| 35.55 | 0.67715 | 3.668 | 0.057704 | 1.2227
| 35.5333 | 0.67001 | 3.6686 | 0.058146 | 1.2229
| 35.5 | 0.67358 | 3.6687 | 0.058673 | 1.2229
| 35.5167 | 0.68072 | 3.669 | 0.058588 | 1.223
| 35.4833 | 0.67358 | 3.6692 | 0.059115 | 1.2231
| 35.5 | 0.67715 | 3.6696 | 0.05903 | 1.2232
| 35.4667 | 0.67358 | 3.6698 | 0.059557 | 1.2233
| 35.4833 | 0.67715 | 3.6702 | 0.059471 | 1.2234
| 35.45 | 0.67358 | 3.6702 | 0.059999 | 1.2234
| 35.4667 | 0.67358 | 3.6706 | 0.059913 | 1.2235
| 35.4333 | 0.67358 | 3.6707 | 0.060441 | 1.2236
| 35.4167 | 0.67715 | 3.6712 | 0.060883 | 1.2237
| 35.45 | 0.67001 | 3.6712 | 0.060355 | 1.2237
| 35.4 | 0.67358 | 3.6716 | 0.061325 | 1.2239
| 35.3833 | 0.67358 | 3.6723 | 0.061767 | 1.2241
| 35.3667 | 0.67715 | 3.6726 | 0.062209 | 1.2242
| 35.4 | 0.67358 | 3.6728 | 0.06168 | 1.2243
| 35.35 | 0.67715 | 3.6732 | 0.06265 | 1.2244
| 35.3833 | 0.67715 | 3.6733 | 0.062122 | 1.2244
| 35.3667 | 0.67715 | 3.6735 | 0.062564 | 1.2245
| 35.3333 | 0.67715 | 3.6736 | 0.063092 | 1.2245
| 35.35 | 0.67358 | 3.6738 | 0.063006 | 1.2246
| 35.3167 | 0.68072 | 3.6739 | 0.063534 | 1.2246
| 35.3 | 0.67358 | 3.6742 | 0.063976 | 1.2247
| 35.2833 | 0.67358 | 3.6746 | 0.064418 | 1.2249
| 35.2667 | 0.67715 | 3.6749 | 0.06486 | 1.225
| 35.2833 | 0.67715 | 3.6751 | 0.064773 | 1.225
| 35.25 | 0.67358 | 3.6753 | 0.065302 | 1.2251
| 35.2333 | 0.67358 | 3.6757 | 0.065744 | 1.2252
| 35.2167 | 0.67715 | 3.676 | 0.066186 | 1.2253
| 35.2333 | 0.67358 | 3.6761 | 0.066098 | 1.2253
| 35.2 | 0.67715 | 3.6764 | 0.066628 | 1.2255
| 35.1833 | 0.67715 | 3.6767 | 0.06707 | 1.2256
| 35.2 | 0.67358 | 3.6767 | 0.066981 | 1.2256
| 35.1667 | 0.67715 | 3.6768 | 0.067512 | 1.2256
| 35.1833 | 0.67358 | 3.6769 | 0.067423 | 1.2256
| 35.1667 | 0.67358 | 3.6771 | 0.067865 | 1.2257
| 35.15 | 0.67715 | 3.6773 | 0.067954 | 1.2258
| 35.1333 | 0.67715 | 3.6774 | 0.068396 | 1.2258
| 35.1167 | 0.67358 | 3.6775 | 0.068838 | 1.2258
| 35.1 | 0.67715 | 3.6777 | 0.06928 | 1.2259
| 35.1167 | 0.67358 | 3.6778 | 0.06919 | 1.2259
| 35.1 | 0.67715 | 3.678 | 0.069632 | 1.226
| 35.0833 | 0.67358 | 3.678 | 0.069721 | 1.226
| 35.0667 | 0.67358 | 3.6783 | 0.070163 | 1.2261
| 35.0667 | 0.67715 | 3.6785 | 0.070516 | 1.2262
| 35.05 | 0.67358 | 3.6786 | 0.070605 | 1.2262
| 35.0333 | 0.67715 | 3.6786 | 0.071047 | 1.2262
| 35.0167 | 0.67358 | 3.6789 | 0.071489 | 1.2263
| 35.0167 | 0.67715 | 3.6791 | 0.071841 | 1.2264
| 35 | 0.67358 | 3.6792 | 0.071931 | 1.2264
| 34.9833 | 0.67358 | 3.6793 | 0.072724 | 1.2264
| 34.9833 | 0.67715 | 3.6793 | 0.072373 | 1.2264
| 34.9667 | 0.67715 | 3.6795 | 0.072815 | 1.2265
| 34.95 | 0.67358 | 3.6796 | 0.073608 | 1.2265
| 34.95 | 0.67358 | 3.6796 | 0.073257 | 1.2265
| 34.9333 | 0.67358 | 3.6799 | 0.073699 | 1.2266
| 34.9167 | 0.67358 | 3.68 | 0.074492 | 1.2267
| 34.9167 | 0.67715 | 3.6802 | 0.074141 | 1.2267
| 34.9 | 0.67715 | 3.6803 | 0.074583 | 1.2268
| 34.8833 | 0.67358 | 3.6805 | 0.075025 | 1.2268
| 34.8667 | 0.67715 | 3.6806 | 0.075467 | 1.2269
| 34.85 | 0.67358 | 3.6807 | 0.076259 | 1.2269
| 34.8167 | 0.67715 | 3.6808 | 0.077142 | 1.2269
| 34.85 | 0.67358 | 3.6809 | 0.075909 | 1.227
| 34.8333 | 0.67715 | 3.6809 | 0.076351 | 1.227
| 34.8167 | 0.67358 | 3.6811 | 0.076792 | 1.227
| 34.8 | 0.67358 | 3.6812 | 0.077234 | 1.2271
| 34.7667 | 0.67358 | 3.6812 | 0.078467 | 1.2271
| 34.7833 | 0.67715 | 3.6813 | 0.077676 | 1.2271
| 34.7667 | 0.67715 | 3.6814 | 0.078118 | 1.2271
| 34.75 | 0.67358 | 3.6815 | 0.078909 | 1.2272
| 34.75 | 0.67001 | 3.6815 | 0.07856 | 1.2272
| 34.7333 | 0.67715 | 3.6816 | 0.079351 | 1.2272
| 34.7 | 0.67715 | 3.6817 | 0.080235 | 1.2272
| 34.7333 | 0.67715 | 3.6818 | 0.079002 | 1.2273
| 34.7167 | 0.67358 | 3.6819 | 0.079444 | 1.2273
| 34.65 | 0.67358 | 3.682 | 0.08156 | 1.2273
| 34.6667 | 0.67715 | 3.6822 | 0.08077 | 1.2274
| 34.65 | 0.68072 | 3.6823 | 0.081212 | 1.2274
| 34.6333 | 0.67358 | 3.6824 | 0.081654 | 1.2275
| 34.6167 | 0.67358 | 3.6825 | 0.082096 | 1.2275
| 34.5833 | 0.67358 | 3.6826 | 0.083327 | 1.2275
| 34.5667 | 0.67358 | 3.6827 | 0.083769 | 1.2276
| 34.6 | 0.67358 | 3.6828 | 0.082538 | 1.2276
| 34.5 | 0.67358 | 3.6829 | 0.085536 | 1.2276
| 34.5833 | 0.67358 | 3.683 | 0.08298 | 1.2277
| 34.4833 | 0.67715 | 3.6831 | 0.085978 | 1.2277
| 34.5667 | 0.67715 | 3.6831 | 0.083422 | 1.2277
| 34.55 | 0.67358 | 3.6832 | 0.083863 | 1.2277
| 34.5167 | 0.67358 | 3.6833 | 0.084747 | 1.2278
| 34.4833 | 0.67715 | 3.6834 | 0.085631 | 1.2278
| 34.4333 | 0.67715 | 3.6835 | 0.087303 | 1.2278
| 34.4667 | 0.67715 | 3.6835 | 0.086073 | 1.2278
| 34.4 | 0.67715 | 3.6836 | 0.088186 | 1.2279
| 34.35 | 0.67358 | 3.6837 | 0.089512 | 1.2279
| 34.45 | 0.67358 | 3.6838 | 0.086515 | 1.2279
| 34.4333 | 0.67715 | 3.6839 | 0.086957 | 1.228
| 34.3833 | 0.67358 | 3.684 | 0.088283 | 1.228
| 34.3667 | 0.67715 | 3.6841 | 0.088725 | 1.228
| 34.35 | 0.67715 | 3.6841 | 0.089167 | 1.228
| 34.3333 | 0.67358 | 3.6842 | 0.089609 | 1.2281
| 34.3167 | 0.67358 | 3.6843 | 0.090051 | 1.2281
| 34.25 | 0.67715 | 3.6844 | 0.092162 | 1.2281
| 34.3 | 0.67715 | 3.6844 | 0.090493 | 1.2281
| 34.2833 | 0.67715 | 3.6845 | 0.090934 | 1.2282
| 34.2667 | 0.67358 | 3.6847 | 0.091376 | 1.2282
| 34.15 | 0.67358 | 3.6848 | 0.094813 | 1.2283
| 34.2333 | 0.67715 | 3.6849 | 0.09226 | 1.2283
| 34.2 | 0.67001 | 3.685 | 0.093144 | 1.2283
| 34.1833 | 0.67715 | 3.685 | 0.093586 | 1.2284
| 34.05 | 0.67358 | 3.6851 | 0.097464 | 1.2284
| 34.0333 | 0.67715 | 3.6852 | 0.097905 | 1.2284
| 34.1667 | 0.67358 | 3.6853 | 0.094028 | 1.2284
| 34.1333 | 0.67358 | 3.6854 | 0.094912 | 1.2285
| 33.9667 | 0.67715 | 3.6855 | 0.099672 | 1.2285
| 34.0667 | 0.67715 | 3.6856 | 0.09668 | 1.2285
| 34.1 | 0.67358 | 3.6857 | 0.095796 | 1.2286
| 34.0333 | 0.67358 | 3.6857 | 0.097564 | 1.2286
| 34.0167 | 0.67358 | 3.6858 | 0.098005 | 1.2286
| 34 | 0.67715 | 3.686 | 0.098447 | 1.2287
| 33.9833 | 0.67358 | 3.686 | 0.098889 | 1.2287
| 33.9667 | 0.67358 | 3.6861 | 0.099331 | 1.2287
| 33.9333 | 0.67358 | 3.6862 | 0.10022 | 1.2287
| 33.9167 | 0.67715 | 3.6863 | 0.10066 | 1.2288
| 33.7167 | 0.67715 | 3.6864 | 0.1063 | 1.2288
| 33.8667 | 0.67358 | 3.6864 | 0.10198 | 1.2288
| 33.6833 | 0.67715 | 3.6866 | 0.10718 | 1.2289
| 33.85 | 0.67358 | 3.6867 | 0.10242 | 1.2289
| 33.7833 | 0.67715 | 3.6868 | 0.10419 | 1.2289
| 33.75 | 0.67358 | 3.6869 | 0.10508 | 1.229
| 33.7167 | 0.67358 | 3.6871 | 0.10596 | 1.229
| 33.7 | 0.67358 | 3.6872 | 0.1064 | 1.2291
| 33.5 | 0.67358 | 3.6873 | 0.11204 | 1.2291
| 33.6833 | 0.67358 | 3.6873 | 0.10684 | 1.2291
| 33.4167 | 0.67358 | 3.6874 | 0.11425 | 1.2291
| 33.6667 | 0.67715 | 3.6875 | 0.10729 | 1.2292
| 33.6167 | 0.67358 | 3.6876 | 0.10861 | 1.2292
| 33.55 | 0.67358 | 3.6877 | 0.11038 | 1.2292
| 33.5833 | 0.67358 | 3.6877 | 0.1095 | 1.2292
| 33.3667 | 0.67715 | 3.6878 | 0.11558 | 1.2293
| 33.3167 | 0.67358 | 3.6879 | 0.1169 | 1.2293
| 33.5167 | 0.67715 | 3.6879 | 0.11126 | 1.2293
| 33.2833 | 0.67358 | 3.688 | 0.11778 | 1.2293
| 33.45 | 0.67715 | 3.6881 | 0.11303 | 1.2294
| 33.4333 | 0.67358 | 3.6882 | 0.11347 | 1.2294
| 33.4 | 0.67358 | 3.6883 | 0.11436 | 1.2294
| 33.4167 | 0.67715 | 3.6883 | 0.11392 | 1.2294
| 33.2167 | 0.67358 | 3.6884 | 0.11955 | 1.2295
| 33.3667 | 0.67358 | 3.6885 | 0.11524 | 1.2295
| 33.35 | 0.67715 | 3.6886 | 0.11568 | 1.2295
| 33.3167 | 0.67715 | 3.6886 | 0.11657 | 1.2295
| 33.2667 | 0.67715 | 3.6887 | 0.11789 | 1.2296
| 33.3 | 0.67358 | 3.6888 | 0.11701 | 1.2296
| 33.25 | 0.67715 | 3.6889 | 0.11833 | 1.2296
| 33.0833 | 0.67715 | 3.6889 | 0.12309 | 1.2296
| 33.2333 | 0.67715 | 3.689 | 0.11878 | 1.2297
| 33.05 | 0.67358 | 3.6891 | 0.12397 | 1.2297
| 33.2167 | 0.67358 | 3.6892 | 0.11922 | 1.2297
| 33.15 | 0.67715 | 3.6892 | 0.12099 | 1.2297
| 33.1167 | 0.67715 | 3.6893 | 0.12187 | 1.2298
| 33.1333 | 0.67715 | 3.6894 | 0.12143 | 1.2298
| 32.9667 | 0.67358 | 3.6895 | 0.12618 | 1.2298
| 33.1 | 0.67358 | 3.6896 | 0.12231 | 1.2299
| 33.0667 | 0.67358 | 3.6897 | 0.1232 | 1.2299
| 32.9333 | 0.67715 | 3.6898 | 0.12706 | 1.2299
| 32.9167 | 0.67358 | 3.6899 | 0.1275 | 1.23
| 33.05 | 0.67001 | 3.6899 | 0.12364 | 1.23
| 33.0333 | 0.67715 | 3.69 | 0.12408 | 1.23
| 33 | 0.67715 | 3.6901 | 0.12496 | 1.23
| 32.9667 | 0.67715 | 3.6902 | 0.12585 | 1.2301
| 32.95 | 0.67715 | 3.6902 | 0.12629 | 1.2301
| 32.9333 | 0.67715 | 3.6904 | 0.12673 | 1.2301
| 32.9167 | 0.67358 | 3.6905 | 0.12717 | 1.2302
| 32.8833 | 0.67715 | 3.6906 | 0.12806 | 1.2302
| 32.7833 | 0.67358 | 3.6907 | 0.13104 | 1.2302
| 32.8667 | 0.67715 | 3.6908 | 0.1285 | 1.2303
| 32.8333 | 0.67715 | 3.6909 | 0.12938 | 1.2303
| 32.8167 | 0.67715 | 3.6909 | 0.12982 | 1.2303
| 32.75 | 0.67358 | 3.691 | 0.13192 | 1.2303
| 32.8 | 0.67715 | 3.6911 | 0.13027 | 1.2304
| 32.7833 | 0.67715 | 3.6912 | 0.13071 | 1.2304
| 32.7167 | 0.67358 | 3.6912 | 0.13281 | 1.2304
| 32.7667 | 0.67358 | 3.6914 | 0.13115 | 1.2305
| 32.6833 | 0.67358 | 3.6915 | 0.13369 | 1.2305
| 32.75 | 0.67715 | 3.6915 | 0.13159 | 1.2305
| 32.6667 | 0.67715 | 3.6916 | 0.13413 | 1.2305
| 32.7167 | 0.67715 | 3.6917 | 0.13248 | 1.2306
| 32.7 | 0.67358 | 3.6918 | 0.13292 | 1.2306
| 32.6833 | 0.67715 | 3.6918 | 0.13336 | 1.2306
| 32.6333 | 0.67358 | 3.6919 | 0.13501 | 1.2306
| 32.6667 | 0.67358 | 3.692 | 0.1338 | 1.2307
| 32.6167 | 0.67715 | 3.6921 | 0.13546 | 1.2307
| 32.65 | 0.67358 | 3.6922 | 0.13424 | 1.2307
| 32.6333 | 0.67715 | 3.6924 | 0.13469 | 1.2308
| 32.6167 | 0.67001 | 3.6926 | 0.13513 | 1.2309
| 32.6 | 0.67358 | 3.6928 | 0.13557 | 1.2309
| 32.5833 | 0.67715 | 3.693 | 0.13601 | 1.231
| 32.5667 | 0.67358 | 3.6931 | 0.13645 | 1.231
| 32.55 | 0.67358 | 3.6932 | 0.1369 | 1.2311
| 32.5333 | 0.67715 | 3.6934 | 0.13734 | 1.2311
| 32.5167 | 0.67358 | 3.6935 | 0.13778 | 1.2312
| 32.5 | 0.67001 | 3.6937 | 0.13855 | 1.2312
| 32.5 | 0.67358 | 3.6939 | 0.13822 | 1.2313
| 32.4833 | 0.67715 | 3.694 | 0.13899 | 1.2313
| 32.4833 | 0.67358 | 3.6941 | 0.13866 | 1.2314
| 32.4667 | 0.67358 | 3.6943 | 0.13943 | 1.2314
| 32.4667 | 0.67358 | 3.6944 | 0.13911 | 1.2315
| 32.45 | 0.67715 | 3.6946 | 0.13987 | 1.2315
| 32.45 | 0.67358 | 3.6947 | 0.13955 | 1.2316
| 32.4333 | 0.67358 | 3.6949 | 0.13999 | 1.2316
| 32.4333 | 0.67715 | 3.6951 | 0.14032 | 1.2317
| 32.4167 | 0.67715 | 3.6952 | 0.14043 | 1.2317
| 32.4167 | 0.67715 | 3.6953 | 0.14076 | 1.2318
| 32.4 | 0.67358 | 3.6953 | 0.14087 | 1.2318
| 32.4 | 0.67358 | 3.6955 | 0.1412 | 1.2318
| 32.3833 | 0.67358 | 3.6956 | 0.14132 | 1.2319
| 32.3833 | 0.67358 | 3.6958 | 0.14164 | 1.2319
| 32.3667 | 0.67358 | 3.6959 | 0.14176 | 1.232
| 32.3667 | 0.67358 | 3.696 | 0.14208 | 1.232
| 32.35 | 0.67358 | 3.6962 | 0.1422 | 1.2321
| 32.35 | 0.67715 | 3.6963 | 0.14252 | 1.2321
| 32.3333 | 0.67715 | 3.6966 | 0.14264 | 1.2322
| 32.3333 | 0.67358 | 3.6967 | 0.14297 | 1.2322
| 32.3167 | 0.67358 | 3.697 | 0.14308 | 1.2323
| 32.3167 | 0.67358 | 3.6971 | 0.14341 | 1.2324
| 32.3 | 0.67358 | 3.6973 | 0.14353 | 1.2324
| 32.3 | 0.67715 | 3.6975 | 0.14385 | 1.2325
| 32.2833 | 0.67358 | 3.6976 | 0.14397 | 1.2325
| 32.2667 | 0.67715 | 3.6978 | 0.14441 | 1.2326
| 32.2833 | 0.67358 | 3.6979 | 0.14429 | 1.2326
| 32.25 | 0.67715 | 3.6982 | 0.14485 | 1.2327
| 32.2333 | 0.67358 | 3.6984 | 0.14529 | 1.2328
| 32.2167 | 0.67715 | 3.6986 | 0.14573 | 1.2329
| 32.2333 | 0.67358 | 3.6987 | 0.14562 | 1.2329
| 32.2 | 0.67715 | 3.6991 | 0.14618 | 1.233
| 32.1833 | 0.67715 | 3.6994 | 0.14662 | 1.2331
| 32.2 | 0.67715 | 3.6995 | 0.1465 | 1.2332
| 32.1833 | 0.67001 | 3.6997 | 0.14694 | 1.2332
| 32.1667 | 0.67715 | 3.6998 | 0.14706 | 1.2333
| 32.1667 | 0.67358 | 3.7 | 0.14738 | 1.2333
| 32.15 | 0.67358 | 3.7001 | 0.1475 | 1.2334
| 32.1333 | 0.67715 | 3.7004 | 0.14794 | 1.2335
| 32.1167 | 0.67358 | 3.7008 | 0.14839 | 1.2336
| 32.1333 | 0.67715 | 3.7008 | 0.14827 | 1.2336
| 32.1 | 0.67715 | 3.701 | 0.14883 | 1.2337
| 32.1167 | 0.67358 | 3.7011 | 0.14871 | 1.2337
| 32.0833 | 0.67358 | 3.7012 | 0.14927 | 1.2337
| 32.1 | 0.67358 | 3.7015 | 0.14915 | 1.2338
| 32.0667 | 0.67358 | 3.7016 | 0.14971 | 1.2339
| 32.0833 | 0.67358 | 3.7017 | 0.14959 | 1.2339
| 32.0667 | 0.67715 | 3.7018 | 0.15003 | 1.2339
| 32.05 | 0.67358 | 3.702 | 0.15015 | 1.234
| 32.0333 | 0.67715 | 3.7022 | 0.1506 | 1.2341
| 32.0333 | 0.67358 | 3.7025 | 0.15092 | 1.2342
| 32.0167 | 0.67358 | 3.7026 | 0.15104 | 1.2342
| 32.0167 | 0.67001 | 3.7028 | 0.15136 | 1.2343
| 32 | 0.67715 | 3.703 | 0.15148 | 1.2343
| 32 | 0.67715 | 3.7031 | 0.1518 | 1.2344
| 31.9833 | 0.67358 | 3.7033 | 0.15192 | 1.2344
| 31.9833 | 0.67715 | 3.7035 | 0.15224 | 1.2345
| 31.9667 | 0.67358 | 3.7036 | 0.15236 | 1.2345
| 31.95 | 0.67715 | 3.7038 | 0.15281 | 1.2346
| 31.9333 | 0.67715 | 3.704 | 0.15325 | 1.2347
| 31.95 | 0.67358 | 3.7042 | 0.15313 | 1.2347
| 31.9167 | 0.67715 | 3.7043 | 0.15369 | 1.2348
| 31.9167 | 0.67358 | 3.7045 | 0.15401 | 1.2348
| 31.9 | 0.67358 | 3.7047 | 0.15413 | 1.2349
| 31.8833 | 0.67358 | 3.705 | 0.15457 | 1.235
| 31.9 | 0.67358 | 3.705 | 0.15445 | 1.235
| 31.8667 | 0.67358 | 3.7053 | 0.15502 | 1.2351
| 31.8667 | 0.67358 | 3.7054 | 0.15534 | 1.2351
| 31.85 | 0.67715 | 3.7056 | 0.15546 | 1.2352
| 31.8333 | 0.67715 | 3.7057 | 0.1559 | 1.2352
| 31.85 | 0.67358 | 3.7058 | 0.15578 | 1.2353
| 31.8333 | 0.67358 | 3.706 | 0.15622 | 1.2353
| 31.8167 | 0.67358 | 3.7061 | 0.15634 | 1.2354
| 31.8 | 0.67358 | 3.7063 | 0.15678 | 1.2354
| 31.8 | 0.67358 | 3.7065 | 0.1571 | 1.2355
| 31.7833 | 0.67715 | 3.7066 | 0.15723 | 1.2355
| 31.7833 | 0.67358 | 3.7068 | 0.15754 | 1.2356
| 31.7667 | 0.67358 | 3.7069 | 0.15767 | 1.2356
| 31.7667 | 0.67715 | 3.7069 | 0.15799 | 1.2356
| 31.75 | 0.67358 | 3.707 | 0.15811 | 1.2357
| 31.75 | 0.67358 | 3.7073 | 0.15843 | 1.2358
| 31.7333 | 0.67715 | 3.7074 | 0.15855 | 1.2358
| 31.7333 | 0.68072 | 3.7076 | 0.15887 | 1.2359
| 31.7167 | 0.67715 | 3.7076 | 0.15899 | 1.2359
| 31.7167 | 0.67001 | 3.7078 | 0.15931 | 1.2359
| 31.7 | 0.67358 | 3.7079 | 0.15943 | 1.236
| 31.7 | 0.67715 | 3.708 | 0.15975 | 1.236
| 31.6833 | 0.67358 | 3.7083 | 0.15988 | 1.2361
| 31.6833 | 0.67358 | 3.7084 | 0.16019 | 1.2361
| 31.6667 | 0.67358 | 3.7085 | 0.16032 | 1.2362
| 31.65 | 0.67358 | 3.7085 | 0.16076 | 1.2362
| 31.6667 | 0.67358 | 3.7086 | 0.16064 | 1.2362
| 31.65 | 0.67001 | 3.7087 | 0.16108 | 1.2362
| 31.6333 | 0.67358 | 3.7089 | 0.1612 | 1.2363
| 31.6333 | 0.67358 | 3.709 | 0.16152 | 1.2363
| 31.6167 | 0.67358 | 3.7091 | 0.16196 | 1.2364
| 31.6167 | 0.67715 | 3.7092 | 0.16164 | 1.2364
| 31.6 | 0.67715 | 3.7092 | 0.1624 | 1.2364
| 31.6 | 0.67358 | 3.7094 | 0.16209 | 1.2365
| 31.5833 | 0.67358 | 3.7095 | 0.16253 | 1.2365
| 31.5833 | 0.67358 | 3.7097 | 0.16285 | 1.2366
| 31.5667 | 0.67715 | 3.7098 | 0.16297 | 1.2366
| 31.55 | 0.67358 | 3.7102 | 0.16341 | 1.2367
| 31.5333 | 0.67715 | 3.7104 | 0.16385 | 1.2368
| 31.5333 | 0.67715 | 3.7105 | 0.16417 | 1.2368
| 31.5167 | 0.67358 | 3.7106 | 0.1643 | 1.2369
| 31.5167 | 0.67358 | 3.7108 | 0.16461 | 1.2369
| 31.5 | 0.67358 | 3.7109 | 0.16474 | 1.237
| 31.4833 | 0.67715 | 3.711 | 0.16518 | 1.237
| 31.4833 | 0.67358 | 3.7111 | 0.1655 | 1.237
| 31.4667 | 0.67358 | 3.7112 | 0.16562 | 1.2371
| 31.45 | 0.67715 | 3.7114 | 0.16638 | 1.2371
| 31.45 | 0.67358 | 3.7115 | 0.16606 | 1.2371
| 31.4333 | 0.67358 | 3.7117 | 0.16651 | 1.2372
| 31.4167 | 0.67358 | 3.7118 | 0.16695 | 1.2373
| 31.4167 | 0.67715 | 3.7119 | 0.16726 | 1.2373
| 31.4 | 0.67358 | 3.7121 | 0.16739 | 1.2374
| 31.4 | 0.67715 | 3.7122 | 0.1677 | 1.2374
| 31.3833 | 0.67715 | 3.7124 | 0.16815 | 1.2375
| 31.3833 | 0.67358 | 3.7124 | 0.16783 | 1.2375
| 31.3667 | 0.67358 | 3.7126 | 0.16859 | 1.2375
| 31.3667 | 0.67358 | 3.7127 | 0.16827 | 1.2376
| 31.35 | 0.67715 | 3.7128 | 0.16872 | 1.2376
| 31.3333 | 0.67715 | 3.713 | 0.16916 | 1.2377
| 31.3167 | 0.67715 | 3.713 | 0.1696 | 1.2377
| 31.3 | 0.67715 | 3.7131 | 0.17036 | 1.2377
| 31.3 | 0.67715 | 3.7134 | 0.17004 | 1.2378
| 31.2833 | 0.68072 | 3.7135 | 0.1708 | 1.2378
| 31.2833 | 0.67358 | 3.7137 | 0.17048 | 1.2379
| 31.2667 | 0.67358 | 3.7137 | 0.17093 | 1.2379
| 31.2667 | 0.67358 | 3.7138 | 0.17124 | 1.2379
| 31.25 | 0.68072 | 3.714 | 0.17168 | 1.238
| 31.25 | 0.67715 | 3.7141 | 0.17137 | 1.238
| 31.2333 | 0.67358 | 3.7143 | 0.17181 | 1.2381
| 31.2333 | 0.67358 | 3.7144 | 0.17212 | 1.2381
| 31.2167 | 0.67001 | 3.7144 | 0.17256 | 1.2381
| 31.2167 | 0.67715 | 3.7145 | 0.17225 | 1.2382
| 31.2 | 0.67001 | 3.7146 | 0.17301 | 1.2382
| 31.1833 | 0.67358 | 3.7147 | 0.17345 | 1.2382
| 31.1833 | 0.67715 | 3.7148 | 0.17313 | 1.2383
| 31.1667 | 0.67358 | 3.7149 | 0.17389 | 1.2383
| 31.1667 | 0.67358 | 3.715 | 0.17358 | 1.2383
| 31.15 | 0.67715 | 3.7151 | 0.17433 | 1.2384
| 31.15 | 0.67715 | 3.7152 | 0.17402 | 1.2384
| 31.1333 | 0.67358 | 3.7153 | 0.17477 | 1.2384
| 31.1333 | 0.67715 | 3.7155 | 0.17446 | 1.2385
| 31.1167 | 0.67715 | 3.7157 | 0.17522 | 1.2386
| 31.1167 | 0.67715 | 3.7157 | 0.1749 | 1.2386
| 31.1 | 0.67001 | 3.7159 | 0.17534 | 1.2386
| 31.0833 | 0.67358 | 3.7161 | 0.1761 | 1.2387
| 31.0833 | 0.67358 | 3.7162 | 0.17579 | 1.2387
| 31.0667 | 0.67715 | 3.7163 | 0.17623 | 1.2388
| 31.05 | 0.67358 | 3.7164 | 0.17698 | 1.2388
| 31.05 | 0.67358 | 3.7166 | 0.17667 | 1.2389
| 31.0333 | 0.67358 | 3.7168 | 0.17711 | 1.2389
| 31.0167 | 0.67358 | 3.7169 | 0.17755 | 1.239
| 31 | 0.67715 | 3.717 | 0.17831 | 1.239
| 31 | 0.67358 | 3.717 | 0.178 | 1.239
| 30.9833 | 0.67358 | 3.7171 | 0.17875 | 1.239
| 30.9833 | 0.67358 | 3.7172 | 0.17844 | 1.2391
| 30.9667 | 0.67715 | 3.7173 | 0.17888 | 1.2391
| 30.95 | 0.67715 | 3.7176 | 0.17932 | 1.2392
| 30.95 | 0.67358 | 3.7177 | 0.17963 | 1.2392
| 30.9333 | 0.67715 | 3.7179 | 0.17976 | 1.2393
| 30.9333 | 0.67358 | 3.718 | 0.18007 | 1.2393
| 30.9167 | 0.67715 | 3.7181 | 0.18052 | 1.2394
| 30.9167 | 0.67715 | 3.7182 | 0.18021 | 1.2394
| 30.9 | 0.67715 | 3.7182 | 0.18065 | 1.2394
| 30.8833 | 0.67358 | 3.7184 | 0.1814 | 1.2395
| 30.8833 | 0.67715 | 3.7185 | 0.18109 | 1.2395
| 30.8667 | 0.67001 | 3.7186 | 0.18184 | 1.2395
| 30.8667 | 0.67358 | 3.7187 | 0.18153 | 1.2396
| 30.85 | 0.67715 | 3.7189 | 0.18228 | 1.2396
| 30.85 | 0.67715 | 3.7189 | 0.18197 | 1.2396
| 30.8333 | 0.67358 | 3.719 | 0.18273 | 1.2397
| 30.8333 | 0.67358 | 3.7191 | 0.18242 | 1.2397
| 30.8167 | 0.67715 | 3.7192 | 0.18286 | 1.2397
| 30.8167 | 0.67715 | 3.7193 | 0.18317 | 1.2398
| 30.8 | 0.67358 | 3.7194 | 0.1833 | 1.2398
| 30.8 | 0.67358 | 3.7196 | 0.18361 | 1.2399
| 30.7833 | 0.67358 | 3.7198 | 0.18374 | 1.2399
| 30.7667 | 0.67715 | 3.7198 | 0.18418 | 1.2399
| 30.7667 | 0.67358 | 3.7201 | 0.18449 | 1.24
| 30.75 | 0.67358 | 3.7202 | 0.18463 | 1.2401
| 30.7333 | 0.67715 | 3.7205 | 0.18538 | 1.2402
| 30.7333 | 0.67358 | 3.7205 | 0.18507 | 1.2402
| 30.7167 | 0.67358 | 3.7206 | 0.18582 | 1.2402
| 30.7167 | 0.67358 | 3.7207 | 0.18551 | 1.2402
| 30.7 | 0.67358 | 3.7208 | 0.18595 | 1.2403
| 30.6833 | 0.67715 | 3.721 | 0.18639 | 1.2403
| 30.6667 | 0.67715 | 3.7211 | 0.18683 | 1.2404
| 30.6667 | 0.67358 | 3.7211 | 0.18714 | 1.2404
| 30.65 | 0.67715 | 3.7214 | 0.18728 | 1.2405
| 30.65 | 0.67358 | 3.7215 | 0.18758 | 1.2405
| 30.6333 | 0.67001 | 3.7217 | 0.18772 | 1.2406
| 30.6333 | 0.67715 | 3.7218 | 0.18803 | 1.2406
| 30.6167 | 0.67358 | 3.7218 | 0.18816 | 1.2406
| 30.6 | 0.67358 | 3.722 | 0.1886 | 1.2407
| 30.6 | 0.67358 | 3.7222 | 0.18891 | 1.2407
| 30.5833 | 0.67358 | 3.7224 | 0.18904 | 1.2408
| 30.5833 | 0.67358 | 3.7224 | 0.18935 | 1.2408
| 30.5667 | 0.67358 | 3.7227 | 0.18949 | 1.2409
| 30.55 | 0.67358 | 3.7228 | 0.19024 | 1.2409
| 30.55 | 0.67001 | 3.7229 | 0.18993 | 1.241
| 30.5333 | 0.67715 | 3.723 | 0.19037 | 1.241
| 30.5167 | 0.67715 | 3.7232 | 0.19081 | 1.2411
| 30.5167 | 0.67715 | 3.7233 | 0.19112 | 1.2411
| 30.5 | 0.67715 | 3.7234 | 0.19125 | 1.2411
| 30.4833 | 0.67358 | 3.7236 | 0.1917 | 1.2412
| 30.5 | 0.67358 | 3.7237 | 0.19156 | 1.2412
| 30.4667 | 0.67715 | 3.7238 | 0.19214 | 1.2413
| 30.45 | 0.67358 | 3.724 | 0.19258 | 1.2413
| 30.4667 | 0.67715 | 3.7243 | 0.19244 | 1.2414
| 30.4333 | 0.67715 | 3.7243 | 0.19302 | 1.2414
| 30.45 | 0.67358 | 3.7245 | 0.19289 | 1.2415
| 30.4167 | 0.67358 | 3.7246 | 0.19346 | 1.2415
| 30.4333 | 0.67358 | 3.7247 | 0.19333 | 1.2415
| 30.4167 | 0.67715 | 3.7248 | 0.19377 | 1.2416
| 30.4 | 0.67715 | 3.725 | 0.19391 | 1.2417
| 30.3833 | 0.67358 | 3.7251 | 0.19435 | 1.2417
| 30.3833 | 0.67715 | 3.7253 | 0.19465 | 1.2418
| 30.3667 | 0.67715 | 3.7254 | 0.19479 | 1.2418
| 30.3667 | 0.67358 | 3.7255 | 0.19509 | 1.2418
| 30.3333 | 0.67358 | 3.7256 | 0.19567 | 1.2419
| 30.3167 | 0.67358 | 3.726 | 0.19612 | 1.242
| 30.3 | 0.67358 | 3.7262 | 0.19656 | 1.2421
| 30.3 | 0.67358 | 3.7265 | 0.19686 | 1.2422
| 30.2833 | 0.67715 | 3.7266 | 0.197 | 1.2422
| 30.2667 | 0.67358 | 3.7267 | 0.19744 | 1.2422
| 30.25 | 0.67358 | 3.7269 | 0.19788 | 1.2423
| 30.2667 | 0.67358 | 3.727 | 0.19775 | 1.2423
| 30.25 | 0.67358 | 3.7272 | 0.19819 | 1.2424
| 30.2333 | 0.67001 | 3.7273 | 0.19833 | 1.2424
| 30.2167 | 0.67715 | 3.7275 | 0.19877 | 1.2425
| 30.2 | 0.67358 | 3.7275 | 0.19921 | 1.2425
| 30.2167 | 0.67715 | 3.7276 | 0.19907 | 1.2425
| 30.1833 | 0.67715 | 3.7278 | 0.19965 | 1.2426
| 30.2 | 0.67358 | 3.7279 | 0.19951 | 1.2426
| 30.1667 | 0.67358 | 3.7281 | 0.20009 | 1.2427
| 30.15 | 0.67358 | 3.7282 | 0.20053 | 1.2427
| 30.1833 | 0.67715 | 3.7283 | 0.19995 | 1.2428
| 30.1667 | 0.67358 | 3.7285 | 0.2004 | 1.2428
| 30.1333 | 0.67715 | 3.7285 | 0.20098 | 1.2428
| 30.1167 | 0.67715 | 3.7288 | 0.20142 | 1.2429
| 30.1333 | 0.67715 | 3.7288 | 0.20128 | 1.2429
| 30.1167 | 0.67715 | 3.729 | 0.20172 | 1.243
| 30.1 | 0.67358 | 3.7291 | 0.20186 | 1.243
| 30.1 | 0.67715 | 3.7292 | 0.20216 | 1.243
| 30.0833 | 0.67715 | 3.7293 | 0.2023 | 1.2431
| 30.0833 | 0.67358 | 3.7294 | 0.2026 | 1.2431
| 30.0667 | 0.67358 | 3.7295 | 0.20274 | 1.2432
| 30.05 | 0.67358 | 3.7296 | 0.20319 | 1.2432
| 30.05 | 0.68072 | 3.7297 | 0.20349 | 1.2432
| 30.0333 | 0.67358 | 3.7298 | 0.20363 | 1.2433
| 30.0167 | 0.67715 | 3.7301 | 0.20407 | 1.2434
| 30 | 0.67715 | 3.7302 | 0.20451 | 1.2434
| 30.0167 | 0.67358 | 3.7303 | 0.20437 | 1.2434
| 29.9833 | 0.67358 | 3.7305 | 0.20495 | 1.2435
| 29.9833 | 0.67358 | 3.7308 | 0.20526 | 1.2436
| 29.9667 | 0.67358 | 3.7308 | 0.2054 | 1.2436
| 29.95 | 0.67358 | 3.731 | 0.20584 | 1.2437
| 29.9333 | 0.67715 | 3.7311 | 0.20628 | 1.2437
| 29.9333 | 0.67715 | 3.7313 | 0.20658 | 1.2438
| 29.9167 | 0.67715 | 3.7314 | 0.20672 | 1.2438
| 29.9 | 0.67001 | 3.7316 | 0.20716 | 1.2439
| 29.9 | 0.67358 | 3.7317 | 0.20746 | 1.2439
| 29.8833 | 0.67358 | 3.7318 | 0.20761 | 1.2439
| 29.8833 | 0.67715 | 3.7319 | 0.20791 | 1.244
| 29.8667 | 0.67715 | 3.732 | 0.20805 | 1.244
| 29.8667 | 0.67358 | 3.7321 | 0.20835 | 1.244
| 29.85 | 0.67358 | 3.7322 | 0.20849 | 1.2441
| 29.8333 | 0.67358 | 3.7323 | 0.20893 | 1.2441
| 29.85 | 0.68072 | 3.7324 | 0.20879 | 1.2441
| 29.8167 | 0.67358 | 3.7326 | 0.20937 | 1.2442
| 29.8333 | 0.67358 | 3.7327 | 0.20923 | 1.2442
| 29.8 | 0.67358 | 3.7329 | 0.20982 | 1.2443
| 29.8 | 0.67358 | 3.733 | 0.21011 | 1.2443
| 29.7833 | 0.67001 | 3.733 | 0.21026 | 1.2443
| 29.7833 | 0.67358 | 3.7331 | 0.21056 | 1.2444
| 29.7667 | 0.67358 | 3.7332 | 0.2107 | 1.2444
| 29.75 | 0.67358 | 3.7334 | 0.21114 | 1.2445
| 29.7333 | 0.67715 | 3.7336 | 0.21188 | 1.2445
| 29.7167 | 0.67358 | 3.7337 | 0.21203 | 1.2446
| 29.7 | 0.67358 | 3.7338 | 0.21247 | 1.2446
| 29.7 | 0.67001 | 3.734 | 0.21277 | 1.2447
| 29.6833 | 0.67358 | 3.7341 | 0.21291 | 1.2447
| 29.6667 | 0.67715 | 3.7343 | 0.21335 | 1.2448
| 29.65 | 0.67715 | 3.7345 | 0.21379 | 1.2448
| 29.6667 | 0.67358 | 3.7346 | 0.21365 | 1.2449
| 29.65 | 0.67358 | 3.7346 | 0.21409 | 1.2449
| 29.6333 | 0.67358 | 3.7347 | 0.21423 | 1.2449
| 29.6333 | 0.68072 | 3.7348 | 0.21453 | 1.2449
| 29.6167 | 0.67358 | 3.7349 | 0.21468 | 1.245
| 29.6167 | 0.67358 | 3.7349 | 0.21497 | 1.245
| 29.6 | 0.67358 | 3.7351 | 0.21512 | 1.245
| 29.5833 | 0.67358 | 3.7353 | 0.21556 | 1.2451
| 29.5667 | 0.67358 | 3.7354 | 0.216 | 1.2451
| 29.55 | 0.67358 | 3.7356 | 0.21644 | 1.2452
| 29.5333 | 0.67715 | 3.7357 | 0.21689 | 1.2452
| 29.5333 | 0.67715 | 3.7358 | 0.21718 | 1.2453
| 29.5 | 0.67715 | 3.7359 | 0.21777 | 1.2453
| 29.5167 | 0.67715 | 3.7359 | 0.21762 | 1.2453
| 29.5 | 0.67358 | 3.7362 | 0.21807 | 1.2454
| 29.4667 | 0.67715 | 3.7363 | 0.21865 | 1.2454
| 29.4833 | 0.67715 | 3.7364 | 0.21851 | 1.2455
| 29.4667 | 0.67358 | 3.7365 | 0.21895 | 1.2455
| 29.45 | 0.67715 | 3.7367 | 0.2191 | 1.2456
| 29.4333 | 0.67715 | 3.7369 | 0.21954 | 1.2456
| 29.4167 | 0.67358 | 3.737 | 0.22028 | 1.2457
| 29.4167 | 0.67358 | 3.7371 | 0.21998 | 1.2457
| 29.4 | 0.67358 | 3.7372 | 0.22042 | 1.2457
| 29.3833 | 0.67358 | 3.7373 | 0.22086 | 1.2458
| 29.3667 | 0.67715 | 3.7374 | 0.22131 | 1.2458
| 29.3667 | 0.67358 | 3.7375 | 0.2216 | 1.2458
| 29.35 | 0.67715 | 3.7376 | 0.22175 | 1.2459
| 29.35 | 0.67358 | 3.7377 | 0.22204 | 1.2459
| 29.3333 | 0.67358 | 3.7378 | 0.22219 | 1.2459
| 29.3333 | 0.67715 | 3.7378 | 0.22248 | 1.2459
| 29.3167 | 0.67358 | 3.7379 | 0.22293 | 1.246
| 29.3167 | 0.67715 | 3.738 | 0.22263 | 1.246
| 29.3 | 0.67358 | 3.7381 | 0.22307 | 1.246
| 29.2833 | 0.67715 | 3.7382 | 0.22381 | 1.2461
| 29.2833 | 0.67715 | 3.7383 | 0.22352 | 1.2461
| 29.25 | 0.67358 | 3.7384 | 0.22469 | 1.2461
| 29.2667 | 0.67358 | 3.7385 | 0.22396 | 1.2462
| 29.2333 | 0.67715 | 3.7386 | 0.22514 | 1.2462
| 29.2333 | 0.67715 | 3.7388 | 0.22484 | 1.2463
| 29.25 | 0.67715 | 3.7388 | 0.2244 | 1.2463
| 29.2167 | 0.67358 | 3.739 | 0.22528 | 1.2463
| 29.2 | 0.67715 | 3.7391 | 0.22573 | 1.2464
| 29.1833 | 0.67358 | 3.7392 | 0.22617 | 1.2464
| 29.1667 | 0.67001 | 3.7393 | 0.22661 | 1.2464
| 29.15 | 0.67358 | 3.7395 | 0.22705 | 1.2465
| 29.1333 | 0.67715 | 3.7395 | 0.22779 | 1.2465
| 29.1333 | 0.67715 | 3.7396 | 0.22749 | 1.2465
| 29.1167 | 0.67715 | 3.7397 | 0.22823 | 1.2466
| 29.1167 | 0.67358 | 3.7398 | 0.22794 | 1.2466
| 29.0833 | 0.67358 | 3.7399 | 0.22911 | 1.2466
| 29.1 | 0.67715 | 3.74 | 0.22838 | 1.2467
| 29.0667 | 0.67715 | 3.7401 | 0.22955 | 1.2467
| 29.0833 | 0.67715 | 3.7401 | 0.22882 | 1.2467
| 29.0667 | 0.67715 | 3.7403 | 0.22926 | 1.2468
| 29.0333 | 0.67358 | 3.7404 | 0.23044 | 1.2468
| 29.05 | 0.67715 | 3.7404 | 0.2297 | 1.2468
| 29.0167 | 0.67715 | 3.7406 | 0.23059 | 1.2469
| 29 | 0.67358 | 3.7407 | 0.23103 | 1.2469
| 29 | 0.67358 | 3.7408 | 0.23132 | 1.2469
| 28.9833 | 0.67358 | 3.7409 | 0.23147 | 1.247
| 28.9667 | 0.67715 | 3.741 | 0.23191 | 1.247
| 28.95 | 0.67358 | 3.741 | 0.23265 | 1.247
| 28.9333 | 0.67715 | 3.7411 | 0.23309 | 1.247
| 28.95 | 0.67358 | 3.7412 | 0.23235 | 1.2471
| 28.9167 | 0.67358 | 3.7413 | 0.23353 | 1.2471
| 28.9333 | 0.67358 | 3.7414 | 0.2328 | 1.2471
| 28.9167 | 0.67358 | 3.7416 | 0.23324 | 1.2472
| 28.8833 | 0.67715 | 3.7417 | 0.23441 | 1.2472
| 28.8833 | 0.67358 | 3.7418 | 0.23412 | 1.2473
| 28.8667 | 0.67358 | 3.7419 | 0.23485 | 1.2473
| 28.85 | 0.67715 | 3.742 | 0.2353 | 1.2473
| 28.85 | 0.67715 | 3.742 | 0.23501 | 1.2473
| 28.8333 | 0.67715 | 3.7421 | 0.23545 | 1.2474
| 28.8167 | 0.67715 | 3.7422 | 0.23589 | 1.2474
| 28.8 | 0.67358 | 3.7423 | 0.23662 | 1.2474
| 28.8 | 0.67715 | 3.7424 | 0.23633 | 1.2474
| 28.7667 | 0.67358 | 3.7424 | 0.2375 | 1.2475
| 28.7833 | 0.67715 | 3.7426 | 0.23677 | 1.2475
| 28.7667 | 0.68072 | 3.7427 | 0.23722 | 1.2476
| 28.7333 | 0.67358 | 3.7427 | 0.23839 | 1.2476
| 28.75 | 0.67358 | 3.7429 | 0.23766 | 1.2476
| 28.7333 | 0.67358 | 3.743 | 0.2381 | 1.2477
| 28.7167 | 0.67358 | 3.7431 | 0.23854 | 1.2477
| 28.7 | 0.67715 | 3.7432 | 0.23927 | 1.2477
| 28.7 | 0.67358 | 3.7433 | 0.23898 | 1.2478
| 28.6667 | 0.67358 | 3.7433 | 0.24016 | 1.2478
| 28.6667 | 0.67358 | 3.7434 | 0.23987 | 1.2478
| 28.65 | 0.67358 | 3.7435 | 0.24031 | 1.2478
| 28.6333 | 0.67358 | 3.7436 | 0.24104 | 1.2479
| 28.6333 | 0.67715 | 3.7437 | 0.24075 | 1.2479
| 28.6 | 0.67358 | 3.7437 | 0.24192 | 1.2479
| 28.6167 | 0.67001 | 3.7439 | 0.24119 | 1.248
| 28.6 | 0.67715 | 3.744 | 0.24164 | 1.248
| 28.5833 | 0.67715 | 3.7443 | 0.24208 | 1.2481
| 28.5667 | 0.67715 | 3.7444 | 0.24252 | 1.2481
| 28.55 | 0.67715 | 3.7445 | 0.24296 | 1.2482
| 28.5167 | 0.67358 | 3.7446 | 0.24413 | 1.2482
| 28.5167 | 0.67715 | 3.7446 | 0.24384 | 1.2482
| 28.5 | 0.67358 | 3.7447 | 0.24457 | 1.2482
| 28.5 | 0.67715 | 3.7448 | 0.24429 | 1.2483
| 28.4667 | 0.67715 | 3.7449 | 0.24517 | 1.2483
| 28.45 | 0.67358 | 3.7449 | 0.2459 | 1.2483
| 28.4333 | 0.67358 | 3.745 | 0.24634 | 1.2483
| 28.45 | 0.67715 | 3.7451 | 0.24561 | 1.2484
| 28.4333 | 0.67715 | 3.7452 | 0.24605 | 1.2484
| 28.4 | 0.67715 | 3.7454 | 0.24722 | 1.2485
| 28.4167 | 0.67715 | 3.7455 | 0.2465 | 1.2485
| 28.4 | 0.67715 | 3.7457 | 0.24694 | 1.2486
| 28.3833 | 0.67358 | 3.7458 | 0.24738 | 1.2486
| 28.3667 | 0.67358 | 3.7459 | 0.24782 | 1.2486
| 28.35 | 0.67358 | 3.7459 | 0.24826 | 1.2486
| 28.3333 | 0.67358 | 3.746 | 0.24871 | 1.2487
| 28.3167 | 0.67715 | 3.7462 | 0.24915 | 1.2487
| 28.2667 | 0.67715 | 3.7462 | 0.25076 | 1.2487
| 28.3 | 0.67715 | 3.7463 | 0.24959 | 1.2488
| 28.2333 | 0.67715 | 3.7465 | 0.25164 | 1.2488
| 28.2833 | 0.67358 | 3.7466 | 0.25003 | 1.2489 | 1.2489
| 28.2667 | 0.67715 | 3.7466 | 0.25047 | 1.2489
| 28.2167 | 0.67358 | 3.7468 | 0.25208 | 1.2489
| 28.25 | 0.67715 | 3.7468 | 0.25092 | 1.249
| 28.2333 | 0.67715 | 3.7469 | 0.25136 | 1.249
| 28.2167 | 0.67715 | 3.747 | 0.2518 | 1.249
| 28.1833 | 0.67358 | 3.7471 | 0.25297 | 1.249
| 28.1833 | 0.67715 | 3.7472 | 0.25268 | 1.2491
| 28.15 | 0.67715 | 3.7473 | 0.25357 | 1.2491
| 28.1667 | 0.67358 | 3.7474 | 0.25313 | 1.2491
| 28.1333 | 0.67358 | 3.7475 | 0.25429 | 1.2492
| 28.1167 | 0.67358 | 3.7475 | 0.25473 | 1.2492
| 28.1333 | 0.67358 | 3.7476 | 0.25401 | 1.2492
| 28.1167 | 0.67715 | 3.7477 | 0.25445 | 1.2492
| 28.1 | 0.67358 | 3.7478 | 0.25518 | 1.2493
| 28.1 | 0.67358 | 3.7478 | 0.25489 | 1.2493
| 28.0667 | 0.67358 | 3.748 | 0.25606 | 1.2493
| 28.05 | 0.67715 | 3.7481 | 0.2565 | 1.2494
| 28.0833 | 0.67715 | 3.7482 | 0.25534 | 1.2494
| 28.0667 | 0.67358 | 3.7483 | 0.25578 | 1.2494
| 28.05 | 0.67358 | 3.7484 | 0.25622 | 1.2495
| 28.0333 | 0.67715 | 3.7485 | 0.25666 | 1.2495
| 28 | 0.67358 | 3.7485 | 0.25783 | 1.2495
| 28.0167 | 0.67358 | 3.7486 | 0.2571 | 1.2495
| 28 | 0.67358 | 3.7487 | 0.25754 | 1.2496
| 27.9833 | 0.67715 | 3.7488 | 0.25799 | 1.2496
| 27.9667 | 0.67715 | 3.7489 | 0.25871 | 1.2496
| 27.9667 | 0.67358 | 3.749 | 0.25843 | 1.2497
| 27.95 | 0.67715 | 3.7491 | 0.25915 | 1.2497
| 27.95 | 0.67358 | 3.7491 | 0.25887 | 1.2497
| 27.9167 | 0.67358 | 3.7493 | 0.26003 | 1.2498
| 27.9333 | 0.67358 | 3.7494 | 0.25931 | 1.2498
| 27.9167 | 0.67358 | 3.7494 | 0.25975 | 1.2498
| 27.8833 | 0.67358 | 3.7495 | 0.26092 | 1.2498
| 27.8833 | 0.67715 | 3.7496 | 0.26064 | 1.2499
| 27.8667 | 0.67358 | 3.7499 | 0.26108 | 1.25
| 27.8333 | 0.67358 | 3.75 | 0.26196 | 1.25
| 27.8333 | 0.67358 | 3.7501 | 0.26224 | 1.25
| 27.8167 | 0.67358 | 3.7502 | 0.26241 | 1.2501
| 27.8 | 0.67715 | 3.7504 | 0.26285 | 1.2501
| 27.7833 | 0.67358 | 3.7504 | 0.26329 | 1.2501
| 27.7833 | 0.67715 | 3.7505 | 0.26357 | 1.2502
| 27.7667 | 0.67358 | 3.7507 | 0.26373 | 1.2502
| 27.75 | 0.67358 | 3.751 | 0.26417 | 1.2503
| 27.7333 | 0.67358 | 3.7511 | 0.26462 | 1.2504
| 27.7167 | 0.67715 | 3.7511 | 0.26534 | 1.2504
| 27.7167 | 0.67358 | 3.7512 | 0.26506 | 1.2504
| 27.7 | 0.67715 | 3.7513 | 0.26578 | 1.2504
| 27.7 | 0.67715 | 3.7514 | 0.2655 | 1.2505
| 27.6833 | 0.67715 | 3.7515 | 0.26622 | 1.2505
| 27.6833 | 0.67715 | 3.7516 | 0.26594 | 1.2505
| 27.6667 | 0.67358 | 3.7517 | 0.26638 | 1.2506
| 27.6333 | 0.67358 | 3.7518 | 0.26727 | 1.2506
| 27.6167 | 0.67715 | 3.7519 | 0.26771 | 1.2506
| 27.6167 | 0.67358 | 3.752 | 0.26799 | 1.2507
| 27.6 | 0.67715 | 3.752 | 0.26815 | 1.2507
| 27.6 | 0.67358 | 3.7521 | 0.26843 | 1.2507
| 27.5833 | 0.67358 | 3.7523 | 0.26859 | 1.2508
| 27.5667 | 0.67358 | 3.7524 | 0.26904 | 1.2508
| 27.55 | 0.67358 | 3.7526 | 0.26975 | 1.2509
| 27.55 | 0.67715 | 3.7527 | 0.26948 | 1.2509
| 27.5333 | 0.67358 | 3.7529 | 0.26992 | 1.251
| 27.5 | 0.67715 | 3.753 | 0.2708 | 1.251
| 27.4833 | 0.67715 | 3.7532 | 0.27124 | 1.2511
| 27.4667 | 0.67358 | 3.7535 | 0.27169 | 1.2512
| 27.45 | 0.67001 | 3.7536 | 0.27213 | 1.2512
| 27.45 | 0.67358 | 3.7536 | 0.2724 | 1.2512
| 27.4333 | 0.67715 | 3.7537 | 0.27257 | 1.2512
| 27.4167 | 0.67715 | 3.7539 | 0.27329 | 1.2513
| 27.4 | 0.67358 | 3.754 | 0.27373 | 1.2513
| 27.4 | 0.67001 | 3.754 | 0.27345 | 1.2513
| 27.3833 | 0.67001 | 3.7542 | 0.2739 | 1.2514
| 27.3667 | 0.67358 | 3.7544 | 0.27461 | 1.2515
| 27.3667 | 0.67715 | 3.7545 | 0.27434 | 1.2515
| 27.35 | 0.67715 | 3.7546 | 0.27478 | 1.2515
| 27.3167 | 0.67358 | 3.7548 | 0.27566 | 1.2516
| 27.3 | 0.67715 | 3.7549 | 0.27611 | 1.2516
| 27.3 | 0.67358 | 3.7551 | 0.27638 | 1.2517
| 27.2833 | 0.67001 | 3.7553 | 0.27655 | 1.2518
| 27.2667 | 0.67715 | 3.7554 | 0.27699 | 1.2518
| 27.25 | 0.67358 | 3.7555 | 0.27771 | 1.2518
| 27.25 | 0.67715 | 3.7556 | 0.27743 | 1.2519
| 27.2333 | 0.67001 | 3.7557 | 0.27787 | 1.2519
| 27.2167 | 0.67358 | 3.7559 | 0.27859 | 1.252
| 27.2167 | 0.67358 | 3.7559 | 0.27832 | 1.252
| 27.2 | 0.67358 | 3.756 | 0.27903 | 1.252
| 27.2 | 0.67358 | 3.7562 | 0.27876 | 1.2521
| 27.1833 | 0.67358 | 3.7562 | 0.27947 | 1.2521
| 27.1833 | 0.67715 | 3.7563 | 0.2792 | 1.2521
| 27.1667 | 0.67715 | 3.7564 | 0.27964 | 1.2521
| 27.15 | 0.67358 | 3.7565 | 0.28008 | 1.2522
| 27.15 | 0.67715 | 3.7566 | 0.28036 | 1.2522
| 27.1333 | 0.67001 | 3.7567 | 0.28053 | 1.2522
| 27.1167 | 0.67715 | 3.7569 | 0.28097 | 1.2523
| 27.1 | 0.67715 | 3.7569 | 0.28141 | 1.2523
| 27.0833 | 0.67358 | 3.7572 | 0.28185 | 1.2524
| 27.0667 | 0.67715 | 3.7572 | 0.28229 | 1.2524
| 27.0667 | 0.67715 | 3.7574 | 0.28257 | 1.2525
| 27.05 | 0.67715 | 3.7575 | 0.28274 | 1.2525
| 27.0333 | 0.67358 | 3.7577 | 0.28318 | 1.2526
| 27.0333 | 0.67358 | 3.7578 | 0.28345 | 1.2526
| 27.0167 | 0.67715 | 3.7578 | 0.28389 | 1.2526
| 27.0167 | 0.67358 | 3.758 | 0.28362 | 1.2527
| 27 | 0.67358 | 3.7581 | 0.28406 | 1.2527
| 26.9833 | 0.67358 | 3.7582 | 0.2845 | 1.2527
| 26.9667 | 0.67715 | 3.7584 | 0.28494 | 1.2528
| 26.9333 | 0.67001 | 3.7585 | 0.28583 | 1.2528
| 26.9333 | 0.67358 | 3.7587 | 0.2861 | 1.2529
| 26.9167 | 0.67358 | 3.7588 | 0.28627 | 1.2529
| 26.9 | 0.67358 | 3.7589 | 0.28671 | 1.253
| 26.9 | 0.67358 | 3.759 | 0.28698 | 1.253
| 26.8833 | 0.67715 | 3.7591 | 0.28742 | 1.253
| 26.8833 | 0.67715 | 3.7591 | 0.28715 | 1.253
| 26.8667 | 0.67358 | 3.7593 | 0.2876 | 1.2531
| 26.85 | 0.67358 | 3.7595 | 0.28831 | 1.2532
| 26.85 | 0.67358 | 3.7596 | 0.28804 | 1.2532
| 26.8333 | 0.67358 | 3.7597 | 0.28875 | 1.2532
| 26.8167 | 0.67715 | 3.7597 | 0.28892 | 1.2532
| 26.8 | 0.67715 | 3.76 | 0.28936 | 1.2533
| 26.7833 | 0.67715 | 3.7601 | 0.28981 | 1.2533
| 26.7833 | 0.67715 | 3.7601 | 0.29008 | 1.2534
| 26.7667 | 0.67358 | 3.7602 | 0.29025 | 1.2534
| 26.75 | 0.67358 | 3.7604 | 0.29069 | 1.2535
| 26.7333 | 0.67358 | 3.7606 | 0.29113 | 1.2535
| 26.7 | 0.67358 | 3.7607 | 0.29228 | 1.2536
| 26.7167 | 0.67001 | 3.7608 | 0.29157 | 1.2536
| 26.6833 | 0.67358 | 3.7609 | 0.29273 | 1.2536
| 26.7 | 0.67358 | 3.761 | 0.29202 | 1.2537
| 26.6833 | 0.67358 | 3.7611 | 0.29246 | 1.2537
| 26.6667 | 0.67715 | 3.7612 | 0.2929 | 1.2537
| 26.65 | 0.67715 | 3.7614 | 0.29334 | 1.2538
| 26.6333 | 0.67358 | 3.7615 | 0.29378 | 1.2538
| 26.6333 | 0.67715 | 3.7616 | 0.29405 | 1.2539
| 26.6167 | 0.67358 | 3.7617 | 0.29423 | 1.2539
| 26.6 | 0.67715 | 3.7618 | 0.29467 | 1.2539
| 26.5833 | 0.67715 | 3.762 | 0.29511 | 1.254
| 26.5667 | 0.67358 | 3.7621 | 0.29582 | 1.254
| 26.55 | 0.67358 | 3.7622 | 0.29626 | 1.2541
| 26.5667 | 0.67358 | 3.7623 | 0.29555 | 1.2541
| 26.55 | 0.67715 | 3.7624 | 0.29599 | 1.2541
| 26.5167 | 0.67715 | 3.7625 | 0.29714 | 1.2542
| 26.5333 | 0.67358 | 3.7626 | 0.29644 | 1.2542
| 26.5 | 0.67715 | 3.7626 | 0.29759 | 1.2542
| 26.5167 | 0.67358 | 3.7627 | 0.29688 | 1.2542
| 26.5 | 0.67358 | 3.7628 | 0.29732 | 1.2543
| 26.4833 | 0.67358 | 3.763 | 0.29776 | 1.2543
| 26.45 | 0.67358 | 3.763 | 0.29891 | 1.2543
| 26.4667 | 0.67715 | 3.7631 | 0.2982 | 1.2544
| 26.45 | 0.68072 | 3.7633 | 0.29865 | 1.2544
| 26.4333 | 0.67358 | 3.7633 | 0.29935 | 1.2544
| 26.4333 | 0.67358 | 3.7634 | 0.29909 | 1.2545
| 26.4167 | 0.67358 | 3.7635 | 0.29979 | 1.2545
| 26.4167 | 0.67358 | 3.7636 | 0.29953 | 1.2545
| 26.4 | 0.67715 | 3.7637 | 0.29997 | 1.2546
| 26.3833 | 0.67358 | 3.7638 | 0.30068 | 1.2546
| 26.3833 | 0.67715 | 3.7639 | 0.30041 | 1.2546
| 26.35 | 0.67358 | 3.7639 | 0.30156 | 1.2546
| 26.3667 | 0.67715 | 3.764 | 0.30085 | 1.2547
| 26.3333 | 0.67715 | 3.7642 | 0.302 | 1.2547
| 26.35 | 0.67358 | 3.7643 | 0.3013 | 1.2548
| 26.3333 | 0.67358 | 3.7643 | 0.30174 | 1.2548
| 26.3167 | 0.67358 | 3.7644 | 0.30218 | 1.2548
| 26.3 | 0.67715 | 3.7645 | 0.30289 | 1.2549
| 26.3 | 0.67358 | 3.7646 | 0.30262 | 1.2549
| 26.2833 | 0.67715 | 3.7647 | 0.30306 | 1.2549
| 26.2667 | 0.67715 | 3.7648 | 0.30377 | 1.2549
| 26.2667 | 0.67358 | 3.7649 | 0.30351 | 1.255
| 26.25 | 0.67358 | 3.765 | 0.30421 | 1.255
| 26.2333 | 0.67715 | 3.7651 | 0.30465 | 1.255
| 26.25 | 0.67358 | 3.7652 | 0.30395 | 1.2551
| 26.2 | 0.67715 | 3.7652 | 0.30554 | 1.2551
| 26.2333 | 0.67358 | 3.7653 | 0.30439 | 1.2551
| 26.2167 | 0.67358 | 3.7655 | 0.30483 | 1.2552
| 26.2 | 0.67001 | 3.7656 | 0.30527 | 1.2552
| 26.15 | 0.67358 | 3.7657 | 0.30686 | 1.2552
| 26.1667 | 0.67715 | 3.7658 | 0.30616 | 1.2553
| 26.15 | 0.67358 | 3.7659 | 0.3066 | 1.2553
| 26.1333 | 0.67358 | 3.766 | 0.3073 | 1.2553
| 26.1333 | 0.67358 | 3.7662 | 0.30704 | 1.2554
| 26.1 | 0.67358 | 3.7662 | 0.30819 | 1.2554
| 26.1167 | 0.67715 | 3.7663 | 0.30748 | 1.2554
| 26.1 | 0.67358 | 3.7665 | 0.30793 | 1.2555
| 26.05 | 0.67715 | 3.7666 | 0.30951 | 1.2555
| 26.0833 | 0.67001 | 3.7667 | 0.30837 | 1.2556
| 26.0667 | 0.67358 | 3.7668 | 0.30881 | 1.2556
| 26.0167 | 0.67715 | 3.7668 | 0.3104 | 1.2556
| 26.05 | 0.67715 | 3.7669 | 0.30925 | 1.2556
| 26 | 0.67715 | 3.767 | 0.31084 | 1.2557
| 26.0167 | 0.67715 | 3.7671 | 0.31014 | 1.2557
| 26.0333 | 0.67715 | 3.7672 | 0.30969 | 1.2557
| 26 | 0.67001 | 3.7674 | 0.31058 | 1.2558
| 25.9833 | 0.67358 | 3.7675 | 0.31102 | 1.2558
| 25.95 | 0.67358 | 3.7675 | 0.31216 | 1.2558
| 25.9667 | 0.67358 | 3.7676 | 0.31146 | 1.2559
| 25.95 | 0.67715 | 3.7678 | 0.3119 | 1.2559
| 25.9167 | 0.67358 | 3.7678 | 0.31305 | 1.2559
| 25.9333 | 0.67715 | 3.768 | 0.31235 | 1.256
| 25.9167 | 0.67715 | 3.7681 | 0.31279 | 1.256
| 25.8667 | 0.67358 | 3.7682 | 0.31437 | 1.2561
| 25.9 | 0.67715 | 3.7683 | 0.31323 | 1.2561
| 25.85 | 0.67715 | 3.7684 | 0.31481 | 1.2561
| 25.8333 | 0.67715 | 3.7684 | 0.31526 | 1.2561
| 25.8667 | 0.67715 | 3.7686 | 0.31411 | 1.2562
| 25.8333 | 0.67358 | 3.7687 | 0.315 | 1.2562
| 25.8167 | 0.67715 | 3.7689 | 0.31544 | 1.2563
| 25.8 | 0.67358 | 3.769 | 0.31588 | 1.2563
| 25.7667 | 0.67358 | 3.7691 | 0.31702 | 1.2564
| 25.75 | 0.67715 | 3.7691 | 0.31747 | 1.2564
| 25.7833 | 0.67358 | 3.7692 | 0.31632 | 1.2564
| 25.7667 | 0.67358 | 3.7694 | 0.31676 | 1.2565
| 25.7 | 0.67358 | 3.7695 | 0.31879 | 1.2565
| 25.6833 | 0.67358 | 3.7696 | 0.31923 | 1.2565
| 25.7333 | 0.67715 | 3.7697 | 0.31765 | 1.2566
| 25.75 | 0.67358 | 3.7698 | 0.31721 | 1.2566
| 25.7167 | 0.67358 | 3.7699 | 0.31809 | 1.2566
| 25.65 | 0.67715 | 3.77 | 0.32012 | 1.2567
| 25.7 | 0.67001 | 3.7701 | 0.31853 | 1.2567
| 25.6833 | 0.67358 | 3.7702 | 0.31897 | 1.2567
| 25.6667 | 0.67715 | 3.7703 | 0.31942 | 1.2568
| 25.65 | 0.67358 | 3.7704 | 0.31986 | 1.2568
| 25.5833 | 0.67715 | 3.7705 | 0.32188 | 1.2568
| 25.5667 | 0.67358 | 3.7706 | 0.32232 | 1.2569
| 25.6167 | 0.67358 | 3.7707 | 0.32074 | 1.2569
| 25.6 | 0.67358 | 3.7708 | 0.32118 | 1.2569
| 25.5333 | 0.67358 | 3.7709 | 0.32321 | 1.257
| 25.5833 | 0.67715 | 3.771 | 0.32163 | 1.257
| 25.5167 | 0.67358 | 3.7711 | 0.32365 | 1.257
| 25.5667 | 0.67358 | 3.7712 | 0.32207 | 1.2571
| 25.55 | 0.67358 | 3.7713 | 0.32251 | 1.2571
| 25.5333 | 0.67001 | 3.7713 | 0.32295 | 1.2571
| 25.5167 | 0.67358 | 3.7715 | 0.32339 | 1.2572
| 25.45 | 0.67715 | 3.7716 | 0.32542 | 1.2572
| 25.5 | 0.67715 | 3.7717 | 0.32384 | 1.2572
| 25.4833 | 0.67715 | 3.7718 | 0.32428 | 1.2573
| 25.4167 | 0.67715 | 3.7719 | 0.3263 | 1.2573
| 25.4667 | 0.67358 | 3.7719 | 0.32472 | 1.2573
| 25.4 | 0.67358 | 3.772 | 0.32674 | 1.2573
| 25.3833 | 0.67358 | 3.7721 | 0.32718 | 1.2574
| 25.45 | 0.67358 | 3.7722 | 0.32516 | 1.2574
| 25.4333 | 0.67001 | 3.7723 | 0.3256 | 1.2574
| 25.35 | 0.67715 | 3.7723 | 0.32807 | 1.2574
| 25.4167 | 0.67715 | 3.7724 | 0.32605 | 1.2575
| 25.3333 | 0.67001 | 3.7726 | 0.32851 | 1.2575
| 25.4 | 0.67358 | 3.7726 | 0.32649 | 1.2575
| 25.3833 | 0.67358 | 3.7728 | 0.32693 | 1.2576
| 25.3667 | 0.67358 | 3.7729 | 0.32737 | 1.2576
| 25.35 | 0.67715 | 3.7729 | 0.32781 | 1.2576
| 25.2667 | 0.67358 | 3.773 | 0.33028 | 1.2577
| 25.3333 | 0.67358 | 3.7731 | 0.32825 | 1.2577
| 25.3167 | 0.67715 | 3.7732 | 0.3287 | 1.2577
| 25.3 | 0.67715 | 3.7733 | 0.32914 | 1.2578
| 25.2167 | 0.67358 | 3.7734 | 0.3316 | 1.2578
| 25.2833 | 0.67358 | 3.7736 | 0.32958 | 1.2579
| 25.2667 | 0.67715 | 3.7736 | 0.33002 | 1.2579
| 25.25 | 0.67715 | 3.7739 | 0.33046 | 1.258
| 25.1667 | 0.67358 | 3.7739 | 0.33293 | 1.258
| 25.2333 | 0.67358 | 3.774 | 0.33091 | 1.258
| 25.2167 | 0.67715 | 3.7741 | 0.33135 | 1.258
| 25.1333 | 0.67715 | 3.7742 | 0.33381 | 1.2581
| 25.2 | 0.67715 | 3.7743 | 0.33179 | 1.2581
| 25.1 | 0.67358 | 3.7744 | 0.33469 | 1.2581
| 25.1667 | 0.67715 | 3.7746 | 0.33267 | 1.2582
| 25.15 | 0.67715 | 3.7746 | 0.33312 | 1.2582
| 25.05 | 0.67358 | 3.7748 | 0.33602 | 1.2583
| 25.1167 | 0.67358 | 3.7748 | 0.334 | 1.2583
| 25.1 | 0.67358 | 3.7751 | 0.33444 | 1.2584
| 25 | 0.67358 | 3.7752 | 0.33734 | 1.2584
| 25.0833 | 0.67715 | 3.7752 | 0.33488 | 1.2584
| 25.05 | 0.67715 | 3.7754 | 0.33577 | 1.2585
| 24.9833 | 0.67715 | 3.7755 | 0.33779 | 1.2585
| 25.0333 | 0.67001 | 3.7756 | 0.33621 | 1.2585
| 24.9167 | 0.67715 | 3.7757 | 0.33955 | 1.2586
| 25 | 0.67715 | 3.7758 | 0.33709 | 1.2586
| 24.9833 | 0.67358 | 3.7759 | 0.33754 | 1.2586
| 24.9 | 0.67358 | 3.776 | 0.34 | 1.2587
| 24.9667 | 0.67358 | 3.7761 | 0.33798 | 1.2587
| 24.9333 | 0.68072 | 3.7763 | 0.33886 | 1.2588
| 24.95 | 0.67715 | 3.7764 | 0.33842 | 1.2588
| 24.85 | 0.67358 | 3.7765 | 0.34132 | 1.2588
| 24.8333 | 0.67715 | 3.7766 | 0.34176 | 1.2589
| 24.9167 | 0.67358 | 3.7767 | 0.3393 | 1.2589
| 24.9 | 0.67001 | 3.7768 | 0.33975 | 1.2589
| 24.8833 | 0.67358 | 3.7768 | 0.34019 | 1.2589
| 24.7833 | 0.67358 | 3.7769 | 0.34309 | 1.259
| 24.7667 | 0.67715 | 3.777 | 0.34353 | 1.259
| 24.8667 | 0.67715 | 3.7771 | 0.34063 | 1.259
| 24.85 | 0.67358 | 3.7771 | 0.34107 | 1.259
| 24.8333 | 0.67358 | 3.7773 | 0.34151 | 1.2591
| 24.8167 | 0.67358 | 3.7774 | 0.34195 | 1.2591
| 24.8 | 0.67715 | 3.7775 | 0.3424 | 1.2592
| 24.7833 | 0.67358 | 3.7776 | 0.34284 | 1.2592
| 24.6833 | 0.67715 | 3.7777 | 0.34574 | 1.2592
| 24.7667 | 0.67358 | 3.7778 | 0.34328 | 1.2592
| 24.6667 | 0.67358 | 3.7778 | 0.34618 | 1.2593
| 24.75 | 0.67715 | 3.7779 | 0.34372 | 1.2593
| 24.7167 | 0.67715 | 3.778 | 0.34461 | 1.2593
| 24.7333 | 0.67358 | 3.7781 | 0.34416 | 1.2594
| 24.6333 | 0.67358 | 3.7781 | 0.34706 | 1.2594
| 24.6167 | 0.67358 | 3.7782 | 0.34751 | 1.2594
| 24.6833 | 0.67715 | 3.7783 | 0.34549 | 1.2594
| 24.7 | 0.67358 | 3.7784 | 0.34505 | 1.2595
| 24.5833 | 0.67358 | 3.7785 | 0.34839 | 1.2595
| 24.6667 | 0.67358 | 3.7786 | 0.34593 | 1.2595
| 24.5667 | 0.67715 | 3.7787 | 0.34883 | 1.2596
| 24.65 | 0.67715 | 3.7787 | 0.34637 | 1.2596
| 24.5333 | 0.67358 | 3.7788 | 0.34971 | 1.2596
| 24.6333 | 0.67358 | 3.7789 | 0.34682 | 1.2596
| 24.6167 | 0.67358 | 3.779 | 0.34726 | 1.2597
| 24.6 | 0.67358 | 3.779 | 0.3477 | 1.2597
| 24.5833 | 0.67358 | 3.7791 | 0.34814 | 1.2597
| 24.5667 | 0.67358 | 3.7793 | 0.34858 | 1.2598
| 24.55 | 0.67715 | 3.7794 | 0.34903 | 1.2598
| 24.4333 | 0.67358 | 3.7795 | 0.35236 | 1.2598
| 24.5333 | 0.67715 | 3.7796 | 0.34947 | 1.2599
| 24.5167 | 0.67001 | 3.7797 | 0.34991 | 1.2599
| 24.4167 | 0.67358 | 3.7797 | 0.35281 | 1.2599
| 24.5 | 0.67358 | 3.7798 | 0.35035 | 1.2599
| 24.3833 | 0.67715 | 3.7799 | 0.35369 | 1.26
| 24.4667 | 0.67358 | 3.78 | 0.35124 | 1.26
| 24.45 | 0.67715 | 3.7802 | 0.35168 | 1.2601
| 24.4333 | 0.67358 | 3.7803 | 0.35212 | 1.2601
| 24.3333 | 0.67715 | 3.7804 | 0.35502 | 1.2601
| 24.4167 | 0.67358 | 3.7805 | 0.35256 | 1.2602
| 24.4 | 0.67358 | 3.7806 | 0.353 | 1.2602
| 24.3 | 0.67358 | 3.7807 | 0.3559 | 1.2602
| 24.3833 | 0.67358 | 3.7807 | 0.35345 | 1.2602
| 24.2667 | 0.67715 | 3.7808 | 0.35678 | 1.2603
| 24.25 | 0.67358 | 3.7809 | 0.35722 | 1.2603
| 24.35 | 0.67715 | 3.781 | 0.35433 | 1.2603
| 24.2333 | 0.67715 | 3.781 | 0.35767 | 1.2603
| 24.3333 | 0.67715 | 3.7812 | 0.35477 | 1.2604
| 24.3167 | 0.67358 | 3.7813 | 0.35521 | 1.2604
| 24.2833 | 0.67358 | 3.7814 | 0.3561 | 1.2605
| 24.1667 | 0.67358 | 3.7815 | 0.35943 | 1.2605
| 24.25 | 0.67715 | 3.7816 | 0.35698 | 1.2605
| 24.1333 | 0.68072 | 3.7817 | 0.36032 | 1.2606
| 24.2167 | 0.67358 | 3.7818 | 0.35786 | 1.2606
| 24.1167 | 0.67715 | 3.7819 | 0.36076 | 1.2606
| 24.1 | 0.67358 | 3.782 | 0.3612 | 1.2607
| 24.2 | 0.67715 | 3.782 | 0.35831 | 1.2607
| 24.1833 | 0.67715 | 3.7821 | 0.35875 | 1.2607
| 24.1667 | 0.67358 | 3.7822 | 0.35919 | 1.2607
| 24.15 | 0.67358 | 3.7823 | 0.35963 | 1.2608
| 24.05 | 0.67358 | 3.7824 | 0.36253 | 1.2608
| 24.1167 | 0.67715 | 3.7826 | 0.36052 | 1.2609
| 24.0833 | 0.67358 | 3.7828 | 0.3614 | 1.2609
| 23.9667 | 0.67358 | 3.7829 | 0.36473 | 1.261
| 24.0667 | 0.67358 | 3.783 | 0.36184 | 1.261
| 23.95 | 0.67358 | 3.7831 | 0.36518 | 1.261
| 24.05 | 0.67715 | 3.7832 | 0.36228 | 1.2611
| 24.0333 | 0.67715 | 3.7833 | 0.36273 | 1.2611
| 23.9 | 0.67715 | 3.7834 | 0.3665 | 1.2611
| 24 | 0.67358 | 3.7835 | 0.36361 | 1.2612
| 23.9833 | 0.67358 | 3.7836 | 0.36405 | 1.2612
| 23.8667 | 0.67358 | 3.7837 | 0.36739 | 1.2612
| 23.9667 | 0.67715 | 3.7838 | 0.36449 | 1.2613
| 23.85 | 0.67715 | 3.7839 | 0.36783 | 1.2613
| 23.9333 | 0.67001 | 3.7839 | 0.36538 | 1.2613
| 23.8167 | 0.67358 | 3.784 | 0.36871 | 1.2613
| 23.9167 | 0.67715 | 3.7841 | 0.36582 | 1.2614
| 23.8833 | 0.67358 | 3.7842 | 0.3667 | 1.2614
| 23.9 | 0.67715 | 3.7843 | 0.36626 | 1.2614
| 23.8667 | 0.67715 | 3.7844 | 0.36715 | 1.2615
| 23.85 | 0.67358 | 3.7846 | 0.36759 | 1.2615
| 23.8167 | 0.68072 | 3.7847 | 0.36847 | 1.2616
| 23.8333 | 0.67715 | 3.7848 | 0.36803 | 1.2616
| 23.8 | 0.67358 | 3.7849 | 0.36891 | 1.2616
| 23.7833 | 0.67358 | 3.7849 | 0.36936 | 1.2616
| 23.6667 | 0.67715 | 3.785 | 0.37269 | 1.2617
| 23.7667 | 0.67358 | 3.7851 | 0.3698 | 1.2617
| 23.65 | 0.67358 | 3.7851 | 0.37313 | 1.2617
| 23.75 | 0.67358 | 3.7852 | 0.37024 | 1.2617
| 23.6167 | 0.67715 | 3.7853 | 0.37401 | 1.2618
| 23.7333 | 0.67358 | 3.7854 | 0.37068 | 1.2618
| 23.7167 | 0.67358 | 3.7855 | 0.37112 | 1.2618
| 23.5667 | 0.67001 | 3.7856 | 0.37534 | 1.2619
| 23.6833 | 0.67715 | 3.7857 | 0.37201 | 1.2619
| 23.6667 | 0.67715 | 3.7858 | 0.37245 | 1.2619
| 23.65 | 0.67715 | 3.7859 | 0.37289 | 1.262
| 23.6333 | 0.67358 | 3.7861 | 0.37333 | 1.262
| 23.6167 | 0.67715 | 3.7861 | 0.37377 | 1.262
| 23.4667 | 0.67358 | 3.7863 | 0.37799 | 1.2621
| 23.5833 | 0.67358 | 3.7864 | 0.37466 | 1.2621
| 23.45 | 0.67358 | 3.7864 | 0.37843 | 1.2621
| 23.4333 | 0.67358 | 3.7865 | 0.37887 | 1.2622
| 23.5667 | 0.67358 | 3.7866 | 0.3751 | 1.2622
| 23.55 | 0.67715 | 3.7867 | 0.37554 | 1.2622
| 23.5167 | 0.67715 | 3.7868 | 0.37643 | 1.2623
| 23.3833 | 0.67715 | 3.7869 | 0.3802 | 1.2623
| 23.5 | 0.67358 | 3.787 | 0.37687 | 1.2623
| 23.4833 | 0.67358 | 3.7871 | 0.37731 | 1.2624
| 23.4667 | 0.67358 | 3.7871 | 0.37775 | 1.2624
| 23.3167 | 0.67715 | 3.7872 | 0.38196 | 1.2624
| 23.45 | 0.67358 | 3.7873 | 0.37819 | 1.2624
| 23.3 | 0.67358 | 3.7874 | 0.38241 | 1.2625
| 23.4167 | 0.67358 | 3.7875 | 0.37908 | 1.2625
| 23.4 | 0.67358 | 3.7877 | 0.37952 | 1.2626
| 23.25 | 0.67358 | 3.7877 | 0.38373 | 1.2626
| 23.3833 | 0.67358 | 3.7878 | 0.37996 | 1.2626
| 23.2333 | 0.67358 | 3.7879 | 0.38417 | 1.2626
| 23.3667 | 0.67358 | 3.788 | 0.3804 | 1.2627
| 23.2 | 0.67358 | 3.7881 | 0.38506 | 1.2627
| 23.3333 | 0.67715 | 3.7881 | 0.38129 | 1.2627
| 23.1667 | 0.67715 | 3.7882 | 0.38594 | 1.2627
| 23.3167 | 0.67358 | 3.7883 | 0.38173 | 1.2628
| 23.2833 | 0.67358 | 3.7884 | 0.38261 | 1.2628
| 23.1333 | 0.67358 | 3.7885 | 0.38682 | 1.2628
| 23.2667 | 0.67715 | 3.7886 | 0.38306 | 1.2629
| 23.25 | 0.67715 | 3.7887 | 0.3835 | 1.2629
| 23.1 | 0.67358 | 3.7887 | 0.38771 | 1.2629
| 23.2333 | 0.67715 | 3.7888 | 0.38394 | 1.2629
| 23.2167 | 0.67358 | 3.7889 | 0.38438 | 1.263
| 23.0333 | 0.67358 | 3.789 | 0.38947 | 1.263
| 23.2 | 0.67358 | 3.789 | 0.38482 | 1.263
| 23.1833 | 0.67715 | 3.7891 | 0.38526 | 1.263
| 23.1667 | 0.67715 | 3.7892 | 0.38571 | 1.2631
| 23.1333 | 0.67358 | 3.7894 | 0.38659 | 1.2631
| 22.9833 | 0.67715 | 3.7894 | 0.3908 | 1.2631
| 23.1167 | 0.67358 | 3.7896 | 0.38703 | 1.2632
| 23.1 | 0.67715 | 3.7897 | 0.38747 | 1.2632
| 22.9333 | 0.67715 | 3.7898 | 0.39212 | 1.2633
| 23.0667 | 0.67358 | 3.7899 | 0.38836 | 1.2633
| 23.05 | 0.67358 | 3.79 | 0.3888 | 1.2633
| 22.8833 | 0.67715 | 3.7901 | 0.39345 | 1.2634
| 23.0333 | 0.67358 | 3.7902 | 0.38924 | 1.2634
| 23.0167 | 0.67715 | 3.7903 | 0.38968 | 1.2634
| 23 | 0.67358 | 3.7903 | 0.39013 | 1.2634
| 22.8333 | 0.67358 | 3.7904 | 0.39477 | 1.2635
| 22.8167 | 0.67358 | 3.7905 | 0.39522 | 1.2635
| 22.9833 | 0.67358 | 3.7906 | 0.39057 | 1.2635
| 22.95 | 0.67715 | 3.7906 | 0.39145 | 1.2635
| 22.8 | 0.67358 | 3.7907 | 0.39566 | 1.2636
| 22.7833 | 0.67715 | 3.7908 | 0.3961 | 1.2636
| 22.9333 | 0.67715 | 3.7909 | 0.39189 | 1.2636
| 22.9167 | 0.67358 | 3.791 | 0.39234 | 1.2636
| 22.9 | 0.67358 | 3.791 | 0.39278 | 1.2637
| 22.8833 | 0.67358 | 3.7912 | 0.39322 | 1.2637
| 22.7 | 0.67715 | 3.7913 | 0.39831 | 1.2638
| 22.8333 | 0.67358 | 3.7913 | 0.39455 | 1.2638
| 22.6833 | 0.67358 | 3.7914 | 0.39875 | 1.2638
| 22.6667 | 0.67715 | 3.7916 | 0.39919 | 1.2639
| 22.8167 | 0.67358 | 3.7916 | 0.39499 | 1.2639
| 22.8 | 0.67358 | 3.7917 | 0.39543 | 1.2639
| 22.6333 | 0.67358 | 3.7918 | 0.40008 | 1.2639
| 22.7667 | 0.67358 | 3.7919 | 0.39631 | 1.264
| 22.75 | 0.67358 | 3.792 | 0.39676 | 1.264
| 22.7333 | 0.67358 | 3.7921 | 0.3972 | 1.264
| 22.5667 | 0.67715 | 3.7922 | 0.40184 | 1.2641
| 22.7167 | 0.67358 | 3.7923 | 0.39764 | 1.2641
| 22.7 | 0.67358 | 3.7923 | 0.39808 | 1.2641
| 22.5167 | 0.67715 | 3.7925 | 0.40317 | 1.2642
| 22.6667 | 0.67358 | 3.7926 | 0.39896 | 1.2642
| 22.4833 | 0.67358 | 3.7927 | 0.40405 | 1.2642
| 22.6333 | 0.67001 | 3.7928 | 0.39985 | 1.2643
| 22.6167 | 0.67358 | 3.7929 | 0.40029 | 1.2643
| 22.4667 | 0.67715 | 3.7929 | 0.40449 | 1.2643
| 22.6 | 0.67001 | 3.793 | 0.40073 | 1.2643
| 22.4333 | 0.67715 | 3.7931 | 0.40538 | 1.2644
| 22.4167 | 0.67358 | 3.7932 | 0.40582 | 1.2644
| 22.5667 | 0.67715 | 3.7932 | 0.40162 | 1.2644
| 22.3833 | 0.67358 | 3.7933 | 0.4067 | 1.2644
| 22.55 | 0.67001 | 3.7934 | 0.40206 | 1.2645
| 22.5333 | 0.67715 | 3.7935 | 0.4025 | 1.2645
| 22.35 | 0.67358 | 3.7936 | 0.40759 | 1.2645
| 22.5 | 0.67715 | 3.7936 | 0.40338 | 1.2645
| 22.3167 | 0.67715 | 3.7938 | 0.40847 | 1.2646
| 22.4833 | 0.67358 | 3.7939 | 0.40383 | 1.2646
| 22.45 | 0.67358 | 3.794 | 0.40471 | 1.2647
| 22.4333 | 0.67358 | 3.7941 | 0.40515 | 1.2647
| 22.2667 | 0.67358 | 3.7942 | 0.40979 | 1.2647
| 22.4167 | 0.67358 | 3.7942 | 0.40559 | 1.2647
| 22.4 | 0.67715 | 3.7943 | 0.40604 | 1.2648
| 22.2167 | 0.67358 | 3.7944 | 0.41112 | 1.2648
| 22.3833 | 0.67358 | 3.7945 | 0.40648 | 1.2648
| 22.2 | 0.67358 | 3.7945 | 0.41156 | 1.2648
| 22.3667 | 0.67358 | 3.7946 | 0.40692 | 1.2649
| 22.35 | 0.67715 | 3.7947 | 0.40736 | 1.2649
| 22.1667 | 0.67715 | 3.7948 | 0.41245 | 1.2649
| 22.3333 | 0.67358 | 3.7948 | 0.4078 | 1.2649
| 22.3167 | 0.67358 | 3.7949 | 0.40825 | 1.265
| 22.3 | 0.67358 | 3.7951 | 0.40869 | 1.265
| 22.2833 | 0.67358 | 3.7952 | 0.40913 | 1.2651
| 22.2667 | 0.67001 | 3.7953 | 0.40957 | 1.2651
| 22.25 | 0.67001 | 3.7955 | 0.41001 | 1.2651
| 22.2333 | 0.67358 | 3.7955 | 0.41046 | 1.2652
| 22.2167 | 0.67358 | 3.7956 | 0.4109 | 1.2652
| 22.0333 | 0.67358 | 3.7957 | 0.41598 | 1.2652
| 22.0167 | 0.67358 | 3.7958 | 0.41642 | 1.2653
| 22.2 | 0.67358 | 3.7958 | 0.41134 | 1.2653
| 21.9833 | 0.67715 | 3.7959 | 0.41731 | 1.2653
| 22.1833 | 0.67715 | 3.796 | 0.41178 | 1.2653
| 22.1667 | 0.67358 | 3.7961 | 0.41222 | 1.2654
| 22.15 | 0.67715 | 3.7961 | 0.41266 | 1.2654
| 22.1333 | 0.67715 | 3.7962 | 0.41311 | 1.2654
| 22.1167 | 0.67358 | 3.7963 | 0.41355 | 1.2654
| 22.1 | 0.67358 | 3.7964 | 0.41399 | 1.2655
| 21.9167 | 0.67358 | 3.7964 | 0.41907 | 1.2655
| 22.0833 | 0.67358 | 3.7966 | 0.41443 | 1.2655
| 22.0667 | 0.67358 | 3.7967 | 0.41487 | 1.2656
| 22.05 | 0.67358 | 3.7967 | 0.41532 | 1.2656
| 22.0333 | 0.67001 | 3.7968 | 0.41576 | 1.2656
| 21.85 | 0.67358 | 3.7969 | 0.42084 | 1.2656
| 22.0167 | 0.67715 | 3.797 | 0.4162 | 1.2657
| 22 | 0.67358 | 3.7971 | 0.41664 | 1.2657
| 21.8333 | 0.67358 | 3.7971 | 0.42128 | 1.2657
| 21.9667 | 0.67715 | 3.7972 | 0.41753 | 1.2657
| 21.9833 | 0.67358 | 3.7973 | 0.41708 | 1.2658
| 21.7833 | 0.67358 | 3.7974 | 0.42261 | 1.2658
| 21.9333 | 0.67358 | 3.7975 | 0.41841 | 1.2658
| 21.75 | 0.67715 | 3.7976 | 0.42349 | 1.2659
| 21.9167 | 0.67358 | 3.7977 | 0.41885 | 1.2659
| 21.7333 | 0.67715 | 3.7977 | 0.42393 | 1.2659
| 21.9 | 0.67358 | 3.7978 | 0.41929 | 1.2659
| 21.7 | 0.67715 | 3.7979 | 0.42482 | 1.266
| 21.8667 | 0.67715 | 3.798 | 0.42018 | 1.266
| 21.6833 | 0.67358 | 3.798 | 0.42526 | 1.266
| 21.85 | 0.67358 | 3.7981 | 0.42062 | 1.266
| 21.6667 | 0.67715 | 3.7982 | 0.4257 | 1.2661
| 21.8333 | 0.67358 | 3.7983 | 0.42106 | 1.2661
| 21.8167 | 0.67715 | 3.7984 | 0.4215 | 1.2661
| 21.8 | 0.67358 | 3.7985 | 0.42195 | 1.2662
| 21.7833 | 0.67715 | 3.7986 | 0.42239 | 1.2662
| 21.5833 | 0.67358 | 3.7987 | 0.42791 | 1.2662
| 21.7667 | 0.67358 | 3.7987 | 0.42283 | 1.2662
| 21.5667 | 0.67001 | 3.7988 | 0.42835 | 1.2663
| 21.75 | 0.67358 | 3.7989 | 0.42327 | 1.2663
| 21.7333 | 0.67715 | 3.799 | 0.42371 | 1.2663
| 21.7167 | 0.67358 | 3.7991 | 0.42416 | 1.2664
| 21.5333 | 0.67715 | 3.7992 | 0.42923 | 1.2664
| 21.7 | 0.67715 | 3.7993 | 0.4246 | 1.2664
| 21.6833 | 0.67358 | 3.7993 | 0.42504 | 1.2664
| 21.65 | 0.67358 | 3.7994 | 0.42592 | 1.2665
| 21.6333 | 0.67358 | 3.7995 | 0.42636 | 1.2665
| 21.4667 | 0.67715 | 3.7996 | 0.431 | 1.2665
| 21.45 | 0.67358 | 3.7997 | 0.43144 | 1.2666
| 21.6167 | 0.67358 | 3.7997 | 0.42681 | 1.2666
| 21.6 | 0.67358 | 3.7999 | 0.42725 | 1.2666
| 21.4167 | 0.67358 | 3.7999 | 0.43233 | 1.2667
| 21.3833 | 0.67358 | 3.8 | 0.43321 | 1.2667
| 21.5667 | 0.67358 | 3.8001 | 0.42813 | 1.2667
| 21.3667 | 0.67358 | 3.8003 | 0.43365 | 1.2668
| 21.55 | 0.67715 | 3.8003 | 0.42857 | 1.2668
| 21.5167 | 0.67715 | 3.8004 | 0.42946 | 1.2668
| 21.5 | 0.67358 | 3.8005 | 0.4299 | 1.2668
| 21.4833 | 0.67358 | 3.8006 | 0.43034 | 1.2669
| 21.3 | 0.67715 | 3.8006 | 0.43542 | 1.2669
| 21.2833 | 0.67358 | 3.8008 | 0.43586 | 1.2669
| 21.45 | 0.67358 | 3.8009 | 0.43123 | 1.267
| 21.4333 | 0.67358 | 3.801 | 0.43167 | 1.267
| 21.4167 | 0.67358 | 3.8012 | 0.43211 | 1.2671
| 21.4 | 0.67358 | 3.8013 | 0.43255 | 1.2671
| 21.3833 | 0.67358 | 3.8014 | 0.43299 | 1.2671
| 21.2 | 0.67715 | 3.8015 | 0.43807 | 1.2672
| 21.35 | 0.67358 | 3.8016 | 0.43388 | 1.2672
| 21.3333 | 0.67715 | 3.8017 | 0.43432 | 1.2672
| 21.3167 | 0.67358 | 3.8019 | 0.43476 | 1.2673
| 21.1333 | 0.67715 | 3.8019 | 0.43984 | 1.2673
| 21.1167 | 0.67358 | 3.802 | 0.44028 | 1.2673
| 21.3 | 0.67715 | 3.8021 | 0.4352 | 1.2674
| 21.2833 | 0.67358 | 3.8022 | 0.43565 | 1.2674
| 21.0833 | 0.67358 | 3.8022 | 0.44116 | 1.2674
| 21.2667 | 0.67358 | 3.8023 | 0.43609 | 1.2674
| 21.2333 | 0.67358 | 3.8024 | 0.43697 | 1.2675
| 21.05 | 0.67715 | 3.8025 | 0.44204 | 1.2675
| 21.2167 | 0.67358 | 3.8026 | 0.43741 | 1.2675
| 21.0333 | 0.67358 | 3.8026 | 0.44249 | 1.2675
| 21 | 0.67715 | 3.8027 | 0.44337 | 1.2676
| 21.2 | 0.67715 | 3.8028 | 0.43786 | 1.2676
| 21.1833 | 0.67358 | 3.8028 | 0.4383 | 1.2676
| 21.1667 | 0.67715 | 3.8029 | 0.43874 | 1.2676
| 21.15 | 0.67715 | 3.803 | 0.43918 | 1.2677
| 21.1333 | 0.67358 | 3.8032 | 0.43962 | 1.2677
| 20.95 | 0.67358 | 3.8032 | 0.44469 | 1.2677
| 20.9333 | 0.67358 | 3.8033 | 0.44514 | 1.2678
| 21.1 | 0.67715 | 3.8034 | 0.44051 | 1.2678
| 20.9 | 0.67715 | 3.8035 | 0.44602 | 1.2678
| 21.0667 | 0.67358 | 3.8036 | 0.44139 | 1.2679
| 21.05 | 0.67358 | 3.8038 | 0.44183 | 1.2679
| 20.85 | 0.67358 | 3.8039 | 0.44735 | 1.268
| 21.0333 | 0.67358 | 3.8039 | 0.44227 | 1.268
| 21.0167 | 0.67715 | 3.8041 | 0.44272 | 1.268
| 21 | 0.67715 | 3.8041 | 0.44316 | 1.268
| 20.9833 | 0.67358 | 3.8043 | 0.4436 | 1.2681
| 20.9667 | 0.67358 | 3.8045 | 0.44404 | 1.2682
| 20.95 | 0.67358 | 3.8046 | 0.44448 | 1.2682
| 20.9333 | 0.67715 | 3.8047 | 0.44493 | 1.2682
| 20.9167 | 0.67358 | 3.8048 | 0.44537 | 1.2683
| 20.7 | 0.67715 | 3.8048 | 0.45132 | 1.2683
| 20.9 | 0.67358 | 3.8049 | 0.44581 | 1.2683
| 20.8667 | 0.67358 | 3.805 | 0.44669 | 1.2683
| 20.8833 | 0.67001 | 3.8051 | 0.44625 | 1.2684
| 20.85 | 0.67715 | 3.8052 | 0.44714 | 1.2684
| 20.6667 | 0.67001 | 3.8053 | 0.4522 | 1.2684
| 20.8333 | 0.67715 | 3.8054 | 0.44758 | 1.2685
| 20.8167 | 0.67001 | 3.8055 | 0.44802 | 1.2685
| 20.8 | 0.67715 | 3.8056 | 0.44846 | 1.2685
| 20.6167 | 0.67715 | 3.8057 | 0.45353 | 1.2686
| 20.7833 | 0.67358 | 3.8058 | 0.4489 | 1.2686
| 20.7667 | 0.67358 | 3.8058 | 0.44935 | 1.2686
| 20.5667 | 0.67358 | 3.8059 | 0.45486 | 1.2686
| 20.7333 | 0.67358 | 3.8061 | 0.45023 | 1.2687
| 20.75 | 0.67715 | 3.8061 | 0.44979 | 1.2687
| 20.7167 | 0.67715 | 3.8063 | 0.45067 | 1.2688
| 20.7 | 0.67358 | 3.8064 | 0.45111 | 1.2688
| 20.6833 | 0.67358 | 3.8064 | 0.45156 | 1.2688
| 20.5 | 0.67715 | 3.8065 | 0.45662 | 1.2688
| 20.6667 | 0.67358 | 3.8066 | 0.452 | 1.2689
| 20.4667 | 0.67358 | 3.8067 | 0.45751 | 1.2689
| 20.45 | 0.67358 | 3.8068 | 0.45795 | 1.2689
| 20.65 | 0.67715 | 3.8069 | 0.45244 | 1.269
| 20.6333 | 0.67358 | 3.807 | 0.45288 | 1.269
| 20.6 | 0.67358 | 3.807 | 0.45377 | 1.269
| 20.4167 | 0.67715 | 3.8071 | 0.45883 | 1.269
| 20.5833 | 0.67358 | 3.8072 | 0.45421 | 1.2691
| 20.5667 | 0.67358 | 3.8074 | 0.45465 | 1.2691
| 20.55 | 0.67715 | 3.8074 | 0.45509 | 1.2691
| 20.5333 | 0.67001 | 3.8077 | 0.45553 | 1.2692
| 20.5167 | 0.67358 | 3.8077 | 0.45597 | 1.2692
| 20.3167 | 0.67715 | 3.8078 | 0.46148 | 1.2693
| 20.4833 | 0.67358 | 3.808 | 0.45686 | 1.2693
| 20.2833 | 0.67358 | 3.808 | 0.46237 | 1.2693
| 20.4667 | 0.67715 | 3.8081 | 0.4573 | 1.2694
| 20.45 | 0.67715 | 3.8082 | 0.45774 | 1.2694
| 20.4333 | 0.67358 | 3.8083 | 0.45818 | 1.2694
| 20.4167 | 0.67358 | 3.8083 | 0.45863 | 1.2694
| 20.2333 | 0.67358 | 3.8084 | 0.46369 | 1.2695
| 20.2167 | 0.67358 | 3.8085 | 0.46413 | 1.2695
| 20.4 | 0.67358 | 3.8087 | 0.45907 | 1.2695
| 20.3833 | 0.67358 | 3.8087 | 0.45951 | 1.2696
| 20.1833 | 0.67715 | 3.8088 | 0.46502 | 1.2696
| 20.3667 | 0.67358 | 3.8089 | 0.45995 | 1.2696
| 20.1667 | 0.67715 | 3.809 | 0.46546 | 1.2697
| 20.35 | 0.67358 | 3.8091 | 0.46039 | 1.2697
| 20.1333 | 0.67358 | 3.8092 | 0.46634 | 1.2697
| 20.3333 | 0.67358 | 3.8093 | 0.46084 | 1.2698
| 20.3167 | 0.67358 | 3.8093 | 0.46128 | 1.2698
| 20.3 | 0.67001 | 3.8094 | 0.46172 | 1.2698
| 20.2833 | 0.67358 | 3.8095 | 0.46216 | 1.2698
| 20.2667 | 0.67358 | 3.8096 | 0.4626 | 1.2699
| 20.0667 | 0.67715 | 3.8097 | 0.46811 | 1.2699
| 20.25 | 0.67715 | 3.8098 | 0.46305 | 1.2699
| 20.2333 | 0.67715 | 3.8099 | 0.46349 | 1.27
| 20.05 | 0.67358 | 3.81 | 0.46855 | 1.27
| 20.0333 | 0.68072 | 3.81 | 0.46899 | 1.27
| 20.2167 | 0.67358 | 3.8101 | 0.46393 | 1.27
| 20.0167 | 0.67715 | 3.8102 | 0.46943 | 1.2701
| 20.2 | 0.67715 | 3.8102 | 0.46437 | 1.2701
| 20.1833 | 0.67358 | 3.8103 | 0.46481 | 1.2701
| 20.1667 | 0.67358 | 3.8104 | 0.46526 | 1.2701
| 19.9667 | 0.67358 | 3.8105 | 0.47076 | 1.2702
| 20.15 | 0.67358 | 3.8106 | 0.4657 | 1.2702
| 20.1167 | 0.67001 | 3.8106 | 0.46658 | 1.2702
| 19.9333 | 0.67358 | 3.8107 | 0.47164 | 1.2702
| 20.1 | 0.67715 | 3.8109 | 0.46702 | 1.2703
| 19.9 | 0.67358 | 3.8109 | 0.47253 | 1.2703
| 20.0833 | 0.67715 | 3.8111 | 0.46747 | 1.2704
| 19.85 | 0.67715 | 3.8112 | 0.47385 | 1.2704
| 19.8667 | 0.67715 | 3.8112 | 0.47341 | 1.2704
| 20.0667 | 0.67715 | 3.8113 | 0.46791 | 1.2704
| 20.05 | 0.67358 | 3.8114 | 0.46835 | 1.2705
| 20.0333 | 0.67358 | 3.8115 | 0.46879 | 1.2705
| 20.0167 | 0.67001 | 3.8116 | 0.46923 | 1.2705
| 19.8167 | 0.67715 | 3.8117 | 0.47474 | 1.2706
| 19.9833 | 0.67358 | 3.8118 | 0.47012 | 1.2706
| 20 | 0.67715 | 3.8119 | 0.46967 | 1.2706
| 19.9667 | 0.67358 | 3.8119 | 0.47056 | 1.2706
| 19.95 | 0.67358 | 3.8121 | 0.471 | 1.2707
| 19.9333 | 0.67715 | 3.8122 | 0.47144 | 1.2707
| 19.7333 | 0.67715 | 3.8123 | 0.47694 | 1.2708
| 19.9167 | 0.67715 | 3.8124 | 0.47188 | 1.2708
| 19.9 | 0.67358 | 3.8125 | 0.47233 | 1.2708
| 19.8833 | 0.67358 | 3.8126 | 0.47277 | 1.2709
| 19.8667 | 0.67358 | 3.8127 | 0.47321 | 1.2709
| 19.6667 | 0.67358 | 3.8129 | 0.47871 | 1.271
| 19.85 | 0.67715 | 3.8129 | 0.47365 | 1.271
| 19.8333 | 0.67358 | 3.813 | 0.47409 | 1.271
| 19.8167 | 0.67715 | 3.8131 | 0.47454 | 1.271
| 19.8 | 0.67001 | 3.8132 | 0.47498 | 1.2711
| 19.7833 | 0.67358 | 3.8133 | 0.47542 | 1.2711
| 19.6 | 0.67358 | 3.8134 | 0.48048 | 1.2711
| 19.7667 | 0.67358 | 3.8135 | 0.47586 | 1.2712
| 19.5667 | 0.67358 | 3.8136 | 0.48136 | 1.2712
| 19.75 | 0.67358 | 3.8137 | 0.4763 | 1.2712
| 19.7333 | 0.67715 | 3.8138 | 0.47675 | 1.2713
| 19.55 | 0.67358 | 3.8139 | 0.4818 | 1.2713
| 19.7167 | 0.67358 | 3.814 | 0.47719 | 1.2713
| 19.7 | 0.67358 | 3.8141 | 0.47763 | 1.2714
| 19.5 | 0.67358 | 3.8142 | 0.48313 | 1.2714
| 19.6667 | 0.67715 | 3.8142 | 0.47851 | 1.2714
| 19.65 | 0.67358 | 3.8143 | 0.47896 | 1.2714
| 19.4667 | 0.67358 | 3.8144 | 0.48401 | 1.2715
| 19.6333 | 0.67715 | 3.8145 | 0.4794 | 1.2715
| 19.45 | 0.67358 | 3.8146 | 0.48445 | 1.2715
| 19.6167 | 0.67715 | 3.8147 | 0.47984 | 1.2716
| 19.6 | 0.67358 | 3.8148 | 0.48028 | 1.2716
| 19.4167 | 0.67358 | 3.8148 | 0.48534 | 1.2716
| 19.5667 | 0.67358 | 3.815 | 0.48117 | 1.2717
| 19.55 | 0.67358 | 3.8152 | 0.48161 | 1.2717
| 19.3667 | 0.67358 | 3.8153 | 0.48666 | 1.2718
| 19.5333 | 0.67715 | 3.8154 | 0.48205 | 1.2718
| 19.5167 | 0.67715 | 3.8154 | 0.48249 | 1.2718
| 19.3333 | 0.67715 | 3.8156 | 0.48755 | 1.2719
| 19.5 | 0.67001 | 3.8157 | 0.48293 | 1.2719
| 19.4833 | 0.67715 | 3.8157 | 0.48337 | 1.2719
| 19.4667 | 0.67358 | 3.8159 | 0.48382 | 1.272
| 19.2667 | 0.67715 | 3.8161 | 0.48931 | 1.272
| 19.45 | 0.67358 | 3.8162 | 0.48426 | 1.2721
| 19.2333 | 0.67358 | 3.8163 | 0.4902 | 1.2721
| 19.4333 | 0.67358 | 3.8164 | 0.4847 | 1.2721
| 19.4 | 0.67358 | 3.8164 | 0.48558 | 1.2721
| 19.2167 | 0.67715 | 3.8165 | 0.49064 | 1.2722
| 19.2 | 0.67715 | 3.8166 | 0.49108 | 1.2722
| 19.3833 | 0.67358 | 3.8167 | 0.48603 | 1.2722
| 19.3667 | 0.67715 | 3.8167 | 0.48647 | 1.2722
| 19.1833 | 0.67715 | 3.8168 | 0.49152 | 1.2723
| 19.1667 | 0.67715 | 3.8169 | 0.49196 | 1.2723
| 19.35 | 0.67715 | 3.817 | 0.48691 | 1.2723
| 19.3333 | 0.67358 | 3.817 | 0.48735 | 1.2723
| 19.1333 | 0.67358 | 3.8172 | 0.49285 | 1.2724
| 19.3167 | 0.67715 | 3.8173 | 0.48779 | 1.2724
| 19.3 | 0.67715 | 3.8174 | 0.48824 | 1.2725
| 19.0833 | 0.67358 | 3.8175 | 0.49417 | 1.2725
| 19.2833 | 0.67715 | 3.8176 | 0.48868 | 1.2725
| 19.2667 | 0.67358 | 3.8176 | 0.48912 | 1.2726
| 19.25 | 0.67358 | 3.8177 | 0.48956 | 1.2726
| 19.2333 | 0.67358 | 3.8179 | 0.49 | 1.2726
| 19.2 | 0.67358 | 3.818 | 0.49089 | 1.2727
| 19.2167 | 0.67358 | 3.8181 | 0.49045 | 1.2727
| 19.0167 | 0.67358 | 3.8182 | 0.49594 | 1.2727
| 19.1833 | 0.67358 | 3.8183 | 0.49133 | 1.2728
| 19 | 0.67358 | 3.8183 | 0.49638 | 1.2728
| 18.9667 | 0.67358 | 3.8185 | 0.49727 | 1.2728
| 19.1667 | 0.67715 | 3.8186 | 0.49177 | 1.2729
| 18.9333 | 0.67358 | 3.8187 | 0.49815 | 1.2729
| 19.1333 | 0.67715 | 3.8188 | 0.49266 | 1.2729
| 19.1167 | 0.67358 | 3.819 | 0.4931 | 1.273
| 19.1 | 0.67358 | 3.8191 | 0.49354 | 1.273
| 19.0833 | 0.67715 | 3.8192 | 0.49398 | 1.2731
| 18.8833 | 0.67358 | 3.8193 | 0.49947 | 1.2731
| 19.0667 | 0.67715 | 3.8193 | 0.49442 | 1.2731
| 19.05 | 0.67358 | 3.8194 | 0.49487 | 1.2731
| 19.0333 | 0.67715 | 3.8195 | 0.49531 | 1.2732
| 19.0167 | 0.67358 | 3.8197 | 0.49575 | 1.2732
| 18.8333 | 0.67715 | 3.8198 | 0.5008 | 1.2733
| 18.8167 | 0.67358 | 3.8199 | 0.50124 | 1.2733
| 18.8 | 0.67715 | 3.8199 | 0.50168 | 1.2733
| 18.9833 | 0.67715 | 3.82 | 0.49663 | 1.2733
| 18.9667 | 0.67715 | 3.8202 | 0.49707 | 1.2734
| 18.95 | 0.67358 | 3.8203 | 0.49752 | 1.2734
| 18.9333 | 0.67358 | 3.8203 | 0.49796 | 1.2734
| 18.75 | 0.67715 | 3.8204 | 0.50301 | 1.2735
| 18.9167 | 0.67358 | 3.8205 | 0.4984 | 1.2735
| 18.9 | 0.67715 | 3.8207 | 0.49884 | 1.2736
| 18.7167 | 0.67001 | 3.8208 | 0.50389 | 1.2736
| 18.8833 | 0.67358 | 3.8209 | 0.49928 | 1.2736
| 18.8667 | 0.67715 | 3.8209 | 0.49973 | 1.2736
| 18.6833 | 0.67358 | 3.821 | 0.50478 | 1.2737
| 18.6667 | 0.67715 | 3.8212 | 0.50522 | 1.2737
| 18.8333 | 0.67715 | 3.8212 | 0.50061 | 1.2737
| 18.8167 | 0.67358 | 3.8215 | 0.50105 | 1.2738
| 18.8 | 0.67358 | 3.8216 | 0.50149 | 1.2739
| 18.7833 | 0.67715 | 3.8218 | 0.50194 | 1.2739
| 18.5833 | 0.67358 | 3.8218 | 0.50743 | 1.2739
| 18.7667 | 0.67715 | 3.8219 | 0.50238 | 1.274
| 18.75 | 0.67358 | 3.822 | 0.50282 | 1.274
| 18.7333 | 0.67358 | 3.8221 | 0.50326 | 1.274
| 18.55 | 0.67358 | 3.8222 | 0.50831 | 1.2741
| 18.7167 | 0.67358 | 3.8223 | 0.5037 | 1.2741
| 18.7 | 0.67358 | 3.8225 | 0.50415 | 1.2742
| 18.6833 | 0.67358 | 3.8226 | 0.50459 | 1.2742
| 18.6667 | 0.67715 | 3.8228 | 0.50503 | 1.2743
| 18.65 | 0.67358 | 3.8228 | 0.50547 | 1.2743
| 18.6333 | 0.67715 | 3.8231 | 0.50591 | 1.2744
| 18.6167 | 0.67358 | 3.8232 | 0.50636 | 1.2744
| 18.6 | 0.67715 | 3.8232 | 0.5068 | 1.2744
| 18.4167 | 0.67358 | 3.8234 | 0.51184 | 1.2745
| 18.5667 | 0.67358 | 3.8234 | 0.50768 | 1.2745
| 18.3833 | 0.67358 | 3.8236 | 0.51273 | 1.2745
| 18.3667 | 0.67358 | 3.8237 | 0.51317 | 1.2746
| 18.55 | 0.67358 | 3.8238 | 0.50812 | 1.2746
| 18.5333 | 0.67715 | 3.8239 | 0.50857 | 1.2746
| 18.5167 | 0.67358 | 3.8241 | 0.50901 | 1.2747
| 18.5 | 0.67715 | 3.8241 | 0.50945 | 1.2747
| 18.4833 | 0.67358 | 3.8243 | 0.50989 | 1.2748
| 18.2833 | 0.67001 | 3.8244 | 0.51538 | 1.2748
| 18.4667 | 0.67358 | 3.8244 | 0.51033 | 1.2748
| 18.45 | 0.67358 | 3.8245 | 0.51077 | 1.2748
| 18.2667 | 0.67715 | 3.8246 | 0.51582 | 1.2749
| 18.4333 | 0.67358 | 3.8247 | 0.51122 | 1.2749
| 18.25 | 0.67358 | 3.8247 | 0.51626 | 1.2749
| 18.4167 | 0.67715 | 3.8249 | 0.51166 | 1.275
| 18.4 | 0.67358 | 3.8251 | 0.5121 | 1.275
| 18.2167 | 0.67358 | 3.8251 | 0.51715 | 1.275
| 18.3833 | 0.67358 | 3.8252 | 0.51254 | 1.2751
| 18.1833 | 0.67358 | 3.8253 | 0.51803 | 1.2751
| 18.3667 | 0.67715 | 3.8254 | 0.51298 | 1.2751
| 18.1667 | 0.67715 | 3.8254 | 0.51847 | 1.2751
| 18.35 | 0.67358 | 3.8255 | 0.51343 | 1.2752
| 18.15 | 0.67715 | 3.8256 | 0.51891 | 1.2752
| 18.3333 | 0.67358 | 3.8257 | 0.51387 | 1.2752
| 18.3167 | 0.67358 | 3.8258 | 0.51431 | 1.2753
| 18.2833 | 0.67358 | 3.826 | 0.51519 | 1.2753
| 18.2667 | 0.67358 | 3.8262 | 0.51564 | 1.2754
| 18.25 | 0.67715 | 3.8264 | 0.51608 | 1.2754
| 18.0667 | 0.67358 | 3.8264 | 0.52112 | 1.2755
| 18.05 | 0.67358 | 3.8265 | 0.52156 | 1.2755
| 18.2333 | 0.67358 | 3.8266 | 0.51652 | 1.2755
| 18.2167 | 0.67358 | 3.8267 | 0.51696 | 1.2756
| 18.0167 | 0.67715 | 3.8268 | 0.52245 | 1.2756
| 18.1833 | 0.67358 | 3.8269 | 0.51785 | 1.2756
| 18 | 0.67715 | 3.827 | 0.52289 | 1.2757
| 18.1667 | 0.67358 | 3.8271 | 0.51829 | 1.2757
| 18.15 | 0.67358 | 3.8272 | 0.51873 | 1.2757
| 17.9667 | 0.67358 | 3.8273 | 0.52377 | 1.2758
| 18.1333 | 0.67358 | 3.8273 | 0.51917 | 1.2758
| 18.1167 | 0.67358 | 3.8275 | 0.51961 | 1.2758
| 17.9333 | 0.67358 | 3.8276 | 0.52466 | 1.2759
| 17.9167 | 0.67715 | 3.8278 | 0.5251 | 1.2759
| 18.0833 | 0.67715 | 3.8279 | 0.5205 | 1.276
| 17.9 | 0.67715 | 3.8281 | 0.52554 | 1.276
| 18.0667 | 0.67358 | 3.8282 | 0.52094 | 1.2761
| 18.05 | 0.67715 | 3.8283 | 0.52138 | 1.2761
| 17.8667 | 0.67715 | 3.8284 | 0.52642 | 1.2761
| 18.0333 | 0.67358 | 3.8285 | 0.52182 | 1.2762
| 18.0167 | 0.67001 | 3.8286 | 0.52227 | 1.2762
| 18 | 0.67358 | 3.8286 | 0.52271 | 1.2762
| 17.9833 | 0.67715 | 3.8288 | 0.52315 | 1.2763
| 17.8 | 0.67358 | 3.8289 | 0.52819 | 1.2763
| 17.9667 | 0.67358 | 3.8289 | 0.52359 | 1.2763
| 17.7833 | 0.67715 | 3.829 | 0.52863 | 1.2763
| 17.95 | 0.67358 | 3.8292 | 0.52403 | 1.2764
| 17.9333 | 0.67715 | 3.8293 | 0.52448 | 1.2764
| 17.9167 | 0.67715 | 3.8294 | 0.52492 | 1.2765
| 17.7333 | 0.67358 | 3.8295 | 0.52996 | 1.2765
| 17.9 | 0.67715 | 3.8296 | 0.52536 | 1.2765
| 17.8833 | 0.68072 | 3.8298 | 0.5258 | 1.2766
| 17.8667 | 0.67358 | 3.8299 | 0.52624 | 1.2766
| 17.6833 | 0.67715 | 3.8299 | 0.53128 | 1.2766
| 17.6667 | 0.67715 | 3.83 | 0.53172 | 1.2767
| 17.85 | 0.67358 | 3.8301 | 0.52668 | 1.2767
| 17.8333 | 0.67358 | 3.8302 | 0.52713 | 1.2767
| 17.65 | 0.67358 | 3.8302 | 0.53217 | 1.2767
| 17.8167 | 0.67358 | 3.8303 | 0.52757 | 1.2768
| 17.8 | 0.67358 | 3.8305 | 0.52801 | 1.2768
| 17.6167 | 0.67358 | 3.8306 | 0.53305 | 1.2769
| 17.6 | 0.67715 | 3.8306 | 0.53349 | 1.2769
| 17.7833 | 0.67358 | 3.8307 | 0.52845 | 1.2769
| 17.5833 | 0.67358 | 3.8308 | 0.53393 | 1.277
| 17.7667 | 0.67715 | 3.8309 | 0.52889 | 1.277
| 17.75 | 0.67358 | 3.831 | 0.52934 | 1.277
| 17.7333 | 0.67358 | 3.8312 | 0.52978 | 1.2771
| 17.7167 | 0.67715 | 3.8312 | 0.53022 | 1.2771
| 17.7 | 0.67358 | 3.8315 | 0.53066 | 1.2772
| 17.6833 | 0.67715 | 3.8315 | 0.5311 | 1.2772
| 17.6667 | 0.67358 | 3.8316 | 0.53155 | 1.2772
| 17.65 | 0.67358 | 3.8318 | 0.53199 | 1.2773
| 17.4833 | 0.67715 | 3.8319 | 0.53658 | 1.2773
| 17.4667 | 0.67358 | 3.832 | 0.53702 | 1.2773
| 17.6333 | 0.67358 | 3.8321 | 0.53243 | 1.2774
| 17.45 | 0.67358 | 3.8321 | 0.53747 | 1.2774
| 17.6167 | 0.67358 | 3.8322 | 0.53287 | 1.2774
| 17.4333 | 0.67715 | 3.8324 | 0.53791 | 1.2775
| 17.6 | 0.67358 | 3.8325 | 0.53331 | 1.2775
| 17.4 | 0.67358 | 3.8325 | 0.53879 | 1.2775
| 17.5667 | 0.67358 | 3.8327 | 0.5342 | 1.2776
| 17.5333 | 0.67358 | 3.8329 | 0.53508 | 1.2776
| 17.35 | 0.67358 | 3.833 | 0.54012 | 1.2777
| 17.5 | 0.67715 | 3.8332 | 0.53597 | 1.2777
| 17.3167 | 0.67358 | 3.8334 | 0.541 | 1.2778
| 17.4833 | 0.67358 | 3.8335 | 0.53641 | 1.2778
| 17.4667 | 0.67358 | 3.8335 | 0.53685 | 1.2778
| 17.3 | 0.67715 | 3.8336 | 0.54144 | 1.2779
| 17.2833 | 0.67715 | 3.8338 | 0.54188 | 1.2779
| 17.45 | 0.67001 | 3.8339 | 0.53729 | 1.278
| 17.4333 | 0.67358 | 3.834 | 0.53773 | 1.278
| 17.25 | 0.67358 | 3.8341 | 0.54277 | 1.278
| 17.4167 | 0.67715 | 3.8342 | 0.53818 | 1.2781
| 17.3833 | 0.67358 | 3.8343 | 0.53906 | 1.2781
| 17.2167 | 0.67358 | 3.8344 | 0.54365 | 1.2781
| 17.3667 | 0.67715 | 3.8345 | 0.5395 | 1.2782
| 17.2 | 0.67358 | 3.8346 | 0.54409 | 1.2782
| 17.35 | 0.67358 | 3.8348 | 0.53994 | 1.2783
| 17.1667 | 0.67358 | 3.8349 | 0.54498 | 1.2783
| 17.3333 | 0.67358 | 3.835 | 0.54038 | 1.2783
| 17.3167 | 0.67715 | 3.8352 | 0.54083 | 1.2784
| 17.1167 | 0.67715 | 3.8353 | 0.5463 | 1.2784
| 17.3 | 0.67715 | 3.8354 | 0.54127 | 1.2785
| 17.2833 | 0.67358 | 3.8355 | 0.54171 | 1.2785
| 17.2667 | 0.67358 | 3.8356 | 0.54215 | 1.2785
| 17.0833 | 0.67715 | 3.8357 | 0.54719 | 1.2786
| 17.2333 | 0.67715 | 3.8358 | 0.54304 | 1.2786
| 17.2167 | 0.67715 | 3.836 | 0.54348 | 1.2787
| 17.0167 | 0.67358 | 3.8362 | 0.54895 | 1.2787
| 17.2 | 0.67358 | 3.8363 | 0.54392 | 1.2788
| 17.1833 | 0.67001 | 3.8364 | 0.54436 | 1.2788
| 17.1667 | 0.67715 | 3.8366 | 0.5448 | 1.2789
| 16.9833 | 0.67715 | 3.8367 | 0.54984 | 1.2789
| 17.15 | 0.67358 | 3.8368 | 0.54525 | 1.2789
| 17.1333 | 0.67358 | 3.8369 | 0.54569 | 1.279
| 17.1167 | 0.67358 | 3.837 | 0.54613 | 1.279
| 17.1 | 0.67715 | 3.8372 | 0.54657 | 1.2791
| 17.0833 | 0.67358 | 3.8374 | 0.54701 | 1.2791
| 17.0667 | 0.67358 | 3.8376 | 0.54746 | 1.2792
| 17.05 | 0.67358 | 3.8377 | 0.5479 | 1.2792
| 16.8667 | 0.67358 | 3.8379 | 0.55293 | 1.2793
| 17.0333 | 0.67358 | 3.838 | 0.54834 | 1.2793
| 17.0167 | 0.67715 | 3.8381 | 0.54878 | 1.2794
| 16.8333 | 0.67358 | 3.8382 | 0.55381 | 1.2794
| 17 | 0.67358 | 3.8382 | 0.54922 | 1.2794
| 16.9833 | 0.67358 | 3.8383 | 0.54967 | 1.2794
| 16.9667 | 0.67358 | 3.8384 | 0.55011 | 1.2795
| 16.8167 | 0.67715 | 3.8385 | 0.55425 | 1.2795
| 16.8 | 0.67358 | 3.8386 | 0.5547 | 1.2795
| 16.7833 | 0.67715 | 3.8386 | 0.55514 | 1.2795
| 16.95 | 0.67715 | 3.8387 | 0.55055 | 1.2796
| 16.9333 | 0.67358 | 3.8388 | 0.55099 | 1.2796
| 16.9167 | 0.67358 | 3.8389 | 0.55143 | 1.2796
| 16.75 | 0.67001 | 3.839 | 0.55602 | 1.2797
| 16.7333 | 0.67358 | 3.8392 | 0.55646 | 1.2797
| 16.9 | 0.67715 | 3.8392 | 0.55188 | 1.2797
| 16.7167 | 0.67358 | 3.8393 | 0.5569 | 1.2798
| 16.8833 | 0.67358 | 3.8394 | 0.55232 | 1.2798
| 16.8667 | 0.67358 | 3.8396 | 0.55276 | 1.2799
| 16.85 | 0.67358 | 3.8397 | 0.5532 | 1.2799
| 16.6667 | 0.67715 | 3.8399 | 0.55823 | 1.28
| 16.8333 | 0.67715 | 3.8399 | 0.55364 | 1.28
| 16.8167 | 0.67358 | 3.8401 | 0.55408 | 1.28
| 16.6333 | 0.67358 | 3.8402 | 0.55911 | 1.2801
| 16.8 | 0.67358 | 3.8402 | 0.55453 | 1.2801
| 16.7833 | 0.67358 | 3.8403 | 0.55497 | 1.2801
| 16.7667 | 0.67358 | 3.8405 | 0.55541 | 1.2802
| 16.6 | 0.67715 | 3.8406 | 0.56 | 1.2802
| 16.75 | 0.67358 | 3.8407 | 0.55585 | 1.2802
| 16.7333 | 0.67001 | 3.8409 | 0.55629 | 1.2803
| 16.7167 | 0.67715 | 3.8409 | 0.55674 | 1.2803
| 16.55 | 0.67358 | 3.8411 | 0.56132 | 1.2804
| 16.7 | 0.67358 | 3.8412 | 0.55718 | 1.2804
| 16.5333 | 0.67358 | 3.8413 | 0.56176 | 1.2804
| 16.6833 | 0.67715 | 3.8414 | 0.55762 | 1.2805
| 16.5167 | 0.67715 | 3.8415 | 0.56221 | 1.2805
| 16.6667 | 0.67358 | 3.8416 | 0.55806 | 1.2805
| 16.65 | 0.67358 | 3.8418 | 0.5585 | 1.2806
| 16.6333 | 0.67358 | 3.8419 | 0.55895 | 1.2806
| 16.6167 | 0.67715 | 3.8421 | 0.55939 | 1.2807
| 16.6 | 0.67715 | 3.8422 | 0.55983 | 1.2807
| 16.4333 | 0.67715 | 3.8423 | 0.56441 | 1.2808
| 16.5833 | 0.67715 | 3.8424 | 0.56027 | 1.2808
| 16.4 | 0.67358 | 3.8426 | 0.5653 | 1.2809
| 16.5667 | 0.67358 | 3.8427 | 0.56071 | 1.2809
| 16.55 | 0.67358 | 3.8428 | 0.56116 | 1.2809
| 16.3833 | 0.67715 | 3.8429 | 0.56574 | 1.281
| 16.3667 | 0.67358 | 3.843 | 0.56618 | 1.281
| 16.5333 | 0.67001 | 3.8431 | 0.5616 | 1.281
| 16.5167 | 0.67358 | 3.8431 | 0.56204 | 1.281
| 16.35 | 0.67358 | 3.8432 | 0.56662 | 1.2811
| 16.5 | 0.67715 | 3.8433 | 0.56248 | 1.2811
| 16.4833 | 0.67715 | 3.8434 | 0.56292 | 1.2811
| 16.4667 | 0.67358 | 3.8435 | 0.56337 | 1.2812
| 16.3 | 0.67715 | 3.8437 | 0.56795 | 1.2812
| 16.45 | 0.67715 | 3.8438 | 0.56381 | 1.2813
| 16.4333 | 0.67358 | 3.8439 | 0.56425 | 1.2813
| 16.2667 | 0.67358 | 3.844 | 0.56883 | 1.2813
| 16.25 | 0.67358 | 3.8442 | 0.56927 | 1.2814
| 16.4167 | 0.67715 | 3.8443 | 0.56469 | 1.2814
| 16.2333 | 0.67358 | 3.8444 | 0.56972 | 1.2815
| 16.4 | 0.67358 | 3.8444 | 0.56513 | 1.2815
| 16.2167 | 0.67358 | 3.8445 | 0.57016 | 1.2815
| 16.3833 | 0.67715 | 3.8447 | 0.56558 | 1.2816
| 16.35 | 0.67715 | 3.8448 | 0.56646 | 1.2816
| 16.1833 | 0.67715 | 3.8449 | 0.57104 | 1.2816
| 16.3333 | 0.67358 | 3.845 | 0.5669 | 1.2817
| 16.3167 | 0.67358 | 3.8451 | 0.56734 | 1.2817
| 16.15 | 0.67715 | 3.8453 | 0.57192 | 1.2818
| 16.3 | 0.67715 | 3.8453 | 0.56778 | 1.2818
| 16.1333 | 0.67715 | 3.8455 | 0.57237 | 1.2818
| 16.2833 | 0.67715 | 3.8456 | 0.56823 | 1.2819
| 16.2667 | 0.67358 | 3.8458 | 0.56867 | 1.2819
| 16.0833 | 0.67358 | 3.846 | 0.57369 | 1.282
| 16.25 | 0.67358 | 3.846 | 0.56911 | 1.282
| 16.0667 | 0.67358 | 3.8461 | 0.57413 | 1.282
| 16.2333 | 0.67358 | 3.8462 | 0.56955 | 1.2821
| 16.2167 | 0.67358 | 3.8463 | 0.56999 | 1.2821
| 16.0333 | 0.67715 | 3.8465 | 0.57502 | 1.2822
| 16.2 | 0.67358 | 3.8466 | 0.57044 | 1.2822
| 16.1833 | 0.67358 | 3.8467 | 0.57088 | 1.2822
| 16.1667 | 0.67358 | 3.8469 | 0.57132 | 1.2823
| 16.15 | 0.67358 | 3.847 | 0.57176 | 1.2823
| 15.9833 | 0.67715 | 3.8471 | 0.57634 | 1.2824
| 16.1333 | 0.67358 | 3.8473 | 0.5722 | 1.2824
| 16.1167 | 0.67358 | 3.8476 | 0.57265 | 1.2825
| 15.9333 | 0.67358 | 3.8476 | 0.57767 | 1.2825
| 16.0833 | 0.67715 | 3.8478 | 0.57353 | 1.2826
| 16.0667 | 0.67715 | 3.8479 | 0.57397 | 1.2826
| 15.9 | 0.67715 | 3.848 | 0.57855 | 1.2827
| 16.05 | 0.67715 | 3.8481 | 0.57441 | 1.2827
| 15.8833 | 0.67715 | 3.8482 | 0.57899 | 1.2827
| 16.0333 | 0.67358 | 3.8483 | 0.57486 | 1.2828
| 16.0167 | 0.67358 | 3.8484 | 0.5753 | 1.2828
| 15.85 | 0.67358 | 3.8485 | 0.57988 | 1.2828
| 16 | 0.67715 | 3.8486 | 0.57574 | 1.2829
| 15.8333 | 0.67358 | 3.8488 | 0.58032 | 1.2829
| 15.9833 | 0.67001 | 3.8489 | 0.57618 | 1.283
| 15.9667 | 0.67358 | 3.849 | 0.57662 | 1.283
| 15.8 | 0.67715 | 3.8491 | 0.5812 | 1.283
| 15.95 | 0.67358 | 3.8493 | 0.57707 | 1.2831
| 15.9167 | 0.67358 | 3.8495 | 0.57795 | 1.2832
| 15.7333 | 0.67358 | 3.8497 | 0.58297 | 1.2832
| 15.9 | 0.67715 | 3.8498 | 0.57839 | 1.2833
| 15.8833 | 0.68072 | 3.8498 | 0.57883 | 1.2833
| 15.8667 | 0.67358 | 3.85 | 0.57928 | 1.2833
| 15.7167 | 0.67715 | 3.8501 | 0.58341 | 1.2834
| 15.7 | 0.67715 | 3.8502 | 0.58385 | 1.2834
| 15.85 | 0.67358 | 3.8502 | 0.57972 | 1.2834
| 15.6833 | 0.67715 | 3.8505 | 0.58429 | 1.2835
| 15.8333 | 0.67358 | 3.8505 | 0.58016 | 1.2835
| 15.6667 | 0.67715 | 3.8506 | 0.58474 | 1.2835
| 15.8167 | 0.67358 | 3.8508 | 0.5806 | 1.2836
| 15.6333 | 0.67358 | 3.8509 | 0.58562 | 1.2836
| 15.8 | 0.67358 | 3.851 | 0.58104 | 1.2837
| 15.7833 | 0.67358 | 3.8512 | 0.58148 | 1.2837
| 15.7667 | 0.67358 | 3.8513 | 0.58193 | 1.2838
| 15.75 | 0.67358 | 3.8514 | 0.58237 | 1.2838
| 15.5833 | 0.67358 | 3.8515 | 0.58694 | 1.2838
| 15.7333 | 0.67715 | 3.8517 | 0.58281 | 1.2839
| 15.7167 | 0.67358 | 3.8519 | 0.58325 | 1.284
| 15.5333 | 0.67358 | 3.852 | 0.58827 | 1.284
| 15.7 | 0.67358 | 3.8521 | 0.58369 | 1.284
| 15.6833 | 0.67358 | 3.8522 | 0.58414 | 1.2841
| 15.5167 | 0.67358 | 3.8523 | 0.58871 | 1.2841
| 15.6667 | 0.67715 | 3.8524 | 0.58458 | 1.2841
| 15.5 | 0.67358 | 3.8525 | 0.58915 | 1.2842
| 15.65 | 0.67715 | 3.8526 | 0.58502 | 1.2842
| 15.4833 | 0.67715 | 3.8527 | 0.5896 | 1.2842
| 15.6333 | 0.67715 | 3.8528 | 0.58546 | 1.2843
| 15.6167 | 0.67715 | 3.853 | 0.5859 | 1.2843
| 15.45 | 0.67358 | 3.8531 | 0.59048 | 1.2844
| 15.6 | 0.67358 | 3.8532 | 0.58635 | 1.2844
| 15.5833 | 0.67358 | 3.8533 | 0.58679 | 1.2844
| 15.4167 | 0.68072 | 3.8534 | 0.59136 | 1.2845
| 15.5667 | 0.67358 | 3.8535 | 0.58723 | 1.2845
| 15.4 | 0.67358 | 3.8537 | 0.5918 | 1.2846
| 15.55 | 0.67358 | 3.8538 | 0.58767 | 1.2846
| 15.5333 | 0.67358 | 3.8539 | 0.58811 | 1.2846
| 15.3667 | 0.67715 | 3.854 | 0.59269 | 1.2847
| 15.5167 | 0.67358 | 3.8542 | 0.58856 | 1.2847
| 15.35 | 0.67358 | 3.8543 | 0.59313 | 1.2848
| 15.5 | 0.67715 | 3.8544 | 0.589 | 1.2848
| 15.3333 | 0.67358 | 3.8545 | 0.59357 | 1.2848
| 15.4833 | 0.67358 | 3.8547 | 0.58944 | 1.2849
| 15.3 | 0.67358 | 3.8547 | 0.59445 | 1.2849
| 15.4667 | 0.67358 | 3.8548 | 0.58988 | 1.2849
| 15.45 | 0.67715 | 3.855 | 0.59032 | 1.285
| 15.4333 | 0.67358 | 3.855 | 0.59077 | 1.285
| 15.2833 | 0.67715 | 3.8551 | 0.5949 | 1.285
| 15.2667 | 0.67715 | 3.8552 | 0.59534 | 1.2851
| 15.4167 | 0.67358 | 3.8553 | 0.59121 | 1.2851
| 15.4 | 0.67715 | 3.8554 | 0.59165 | 1.2851
| 15.2333 | 0.67715 | 3.8556 | 0.59622 | 1.2852
| 15.3833 | 0.67715 | 3.8557 | 0.59209 | 1.2852
| 15.2167 | 0.67358 | 3.8558 | 0.59666 | 1.2853
| 15.3667 | 0.67715 | 3.8559 | 0.59253 | 1.2853
| 15.2 | 0.67358 | 3.8561 | 0.59711 | 1.2854
| 15.35 | 0.67358 | 3.8563 | 0.59298 | 1.2854
| 15.3333 | 0.67358 | 3.8563 | 0.59342 | 1.2854
| 15.3167 | 0.67715 | 3.8566 | 0.59386 | 1.2855
| 15.3 | 0.67358 | 3.8566 | 0.5943 | 1.2855
| 15.1333 | 0.67358 | 3.8567 | 0.59887 | 1.2856
| 15.2833 | 0.67358 | 3.8569 | 0.59474 | 1.2856
| 15.2667 | 0.67001 | 3.857 | 0.59518 | 1.2857
| 15.25 | 0.67715 | 3.8573 | 0.59563 | 1.2857
| 15.1 | 0.67358 | 3.8573 | 0.59976 | 1.2858
| 15.2333 | 0.67358 | 3.8575 | 0.59607 | 1.2858
| 15.0667 | 0.67715 | 3.8577 | 0.60064 | 1.2859
| 15.2167 | 0.67715 | 3.8578 | 0.59651 | 1.2859
| 15.05 | 0.67358 | 3.8579 | 0.60108 | 1.286
| 15.2 | 0.67715 | 3.858 | 0.59695 | 1.286
| 15.0333 | 0.67715 | 3.8581 | 0.60152 | 1.286
| 15.1833 | 0.67715 | 3.8582 | 0.59739 | 1.2861
| 15.0167 | 0.67715 | 3.8583 | 0.60196 | 1.2861
| 15.1667 | 0.67715 | 3.8584 | 0.59784 | 1.2861
| 15.15 | 0.67715 | 3.8586 | 0.59828 | 1.2862
| 15.1333 | 0.67358 | 3.8586 | 0.59872 | 1.2862
| 14.9667 | 0.67715 | 3.8587 | 0.60329 | 1.2862
| 15.1167 | 0.67358 | 3.8588 | 0.59916 | 1.2863
| 14.95 | 0.67358 | 3.8589 | 0.60373 | 1.2863
| 15.1 | 0.67358 | 3.8591 | 0.5996 | 1.2864
| 14.9333 | 0.67715 | 3.8593 | 0.60417 | 1.2864
| 15.0833 | 0.67358 | 3.8594 | 0.60005 | 1.2865
| 15.0667 | 0.67358 | 3.8595 | 0.60049 | 1.2865
| 14.9167 | 0.67715 | 3.8596 | 0.60462 | 1.2865
| 15.05 | 0.67358 | 3.8597 | 0.60093 | 1.2866
| 14.9 | 0.67715 | 3.8598 | 0.60506 | 1.2866
| 15.0333 | 0.67358 | 3.8599 | 0.60137 | 1.2866
| 14.8833 | 0.67715 | 3.8601 | 0.6055 | 1.2867
| 15.0167 | 0.67358 | 3.8601 | 0.60181 | 1.2867
| 14.8667 | 0.67358 | 3.8602 | 0.60594 | 1.2867
| 15 | 0.67358 | 3.8603 | 0.60226 | 1.2868
| 14.85 | 0.67358 | 3.8604 | 0.60638 | 1.2868
| 14.9833 | 0.67358 | 3.8605 | 0.6027 | 1.2868
| 14.8333 | 0.67358 | 3.8606 | 0.60682 | 1.2869
| 14.9667 | 0.67358 | 3.8608 | 0.60314 | 1.2869
| 14.8167 | 0.67715 | 3.8608 | 0.60727 | 1.2869
| 14.8 | 0.67715 | 3.861 | 0.60771 | 1.287
| 14.95 | 0.67358 | 3.8611 | 0.60358 | 1.287
| 14.9333 | 0.67715 | 3.8611 | 0.60402 | 1.287
| 14.7833 | 0.67358 | 3.8612 | 0.60815 | 1.2871
| 14.9167 | 0.67715 | 3.8614 | 0.60447 | 1.2871
| 14.7667 | 0.67358 | 3.8615 | 0.60859 | 1.2872
| 14.75 | 0.67715 | 3.8616 | 0.60903 | 1.2872
| 14.9 | 0.67358 | 3.8617 | 0.60491 | 1.2872
| 14.8833 | 0.67715 | 3.8619 | 0.60535 | 1.2873
| 14.8667 | 0.67715 | 3.8621 | 0.60579 | 1.2874
| 14.7167 | 0.67358 | 3.8621 | 0.60992 | 1.2874
| 14.85 | 0.67715 | 3.8622 | 0.60623 | 1.2874
| 14.7 | 0.67358 | 3.8623 | 0.61036 | 1.2874
| 14.8333 | 0.67358 | 3.8624 | 0.60668 | 1.2875
| 14.8167 | 0.67715 | 3.8626 | 0.60712 | 1.2875
| 14.6667 | 0.67358 | 3.8627 | 0.61124 | 1.2876
| 14.8 | 0.67715 | 3.8628 | 0.60756 | 1.2876
| 14.65 | 0.67715 | 3.8629 | 0.61168 | 1.2876
| 14.6333 | 0.67358 | 3.863 | 0.61213 | 1.2877
| 14.7833 | 0.67358 | 3.8631 | 0.608 | 1.2877
| 14.7667 | 0.67358 | 3.8634 | 0.60844 | 1.2878
| 14.6167 | 0.67715 | 3.8635 | 0.61257 | 1.2878
| 14.75 | 0.67715 | 3.8636 | 0.60889 | 1.2879
| 14.6 | 0.67715 | 3.8637 | 0.61301 | 1.2879
| 14.7167 | 0.67358 | 3.8639 | 0.60977 | 1.288
| 14.5667 | 0.67358 | 3.8641 | 0.61389 | 1.288
| 14.7 | 0.67358 | 3.8642 | 0.61021 | 1.2881
| 14.55 | 0.67715 | 3.8643 | 0.61433 | 1.2881
| 14.6833 | 0.67358 | 3.8643 | 0.61065 | 1.2881
| 14.5333 | 0.67358 | 3.8645 | 0.61478 | 1.2882
| 14.6667 | 0.67358 | 3.8646 | 0.61109 | 1.2882
| 14.65 | 0.67715 | 3.8647 | 0.61154 | 1.2882
| 14.5167 | 0.67358 | 3.8647 | 0.61522 | 1.2882
| 14.5 | 0.67358 | 3.865 | 0.61566 | 1.2883
| 14.6333 | 0.67358 | 3.865 | 0.61198 | 1.2883
| 14.6167 | 0.67358 | 3.8652 | 0.61242 | 1.2884
| 14.6 | 0.67358 | 3.8655 | 0.61286 | 1.2885
| 14.5833 | 0.67715 | 3.8656 | 0.6133 | 1.2885
| 14.45 | 0.67715 | 3.8657 | 0.61699 | 1.2886
| 14.4333 | 0.67358 | 3.866 | 0.61743 | 1.2887
| 14.5667 | 0.67715 | 3.866 | 0.61375 | 1.2887
| 14.55 | 0.67358 | 3.8662 | 0.61419 | 1.2887
| 14.5333 | 0.67358 | 3.8662 | 0.61463 | 1.2888
| 14.5167 | 0.67358 | 3.8665 | 0.61507 | 1.2888
| 14.3833 | 0.67358 | 3.8666 | 0.61875 | 1.2889
| 14.5 | 0.67715 | 3.8667 | 0.61551 | 1.2889
| 14.4833 | 0.67715 | 3.867 | 0.61596 | 1.289
| 14.35 | 0.67715 | 3.8671 | 0.61964 | 1.289
| 14.3333 | 0.67715 | 3.8672 | 0.62008 | 1.2891
| 14.4667 | 0.67358 | 3.8672 | 0.6164 | 1.2891
| 14.45 | 0.67358 | 3.8675 | 0.61684 | 1.2892
| 14.3167 | 0.67715 | 3.8676 | 0.62052 | 1.2892
| 14.4333 | 0.67358 | 3.8678 | 0.61728 | 1.2893
| 14.2833 | 0.67715 | 3.868 | 0.6214 | 1.2893
| 14.4167 | 0.67358 | 3.8681 | 0.61772 | 1.2894
| 14.3833 | 0.67358 | 3.8683 | 0.61861 | 1.2894
| 14.3667 | 0.67715 | 3.8685 | 0.61905 | 1.2895
| 14.2333 | 0.67358 | 3.8687 | 0.62273 | 1.2896
| 14.35 | 0.67358 | 3.8689 | 0.61949 | 1.2896
| 14.3333 | 0.67358 | 3.869 | 0.61993 | 1.2897
| 14.2 | 0.67358 | 3.8692 | 0.62361 | 1.2897
| 14.3167 | 0.67715 | 3.8694 | 0.62038 | 1.2898
| 14.3 | 0.67715 | 3.8695 | 0.62082 | 1.2898
| 14.1667 | 0.67358 | 3.8697 | 0.6245 | 1.2899
| 14.2833 | 0.67358 | 3.8698 | 0.62126 | 1.2899
| 14.2667 | 0.67715 | 3.8701 | 0.6217 | 1.29
| 14.1333 | 0.67358 | 3.8702 | 0.62538 | 1.2901
| 14.25 | 0.67358 | 3.8704 | 0.62214 | 1.2901
| 14.2333 | 0.67358 | 3.8704 | 0.62259 | 1.2901
| 14.1 | 0.67358 | 3.8706 | 0.62626 | 1.2902
| 14.2167 | 0.67715 | 3.8707 | 0.62303 | 1.2902
| 14.2 | 0.67358 | 3.8709 | 0.62347 | 1.2903
| 14.1833 | 0.67358 | 3.8711 | 0.62391 | 1.2904
| 14.05 | 0.67358 | 3.8713 | 0.62759 | 1.2904
| 14.1667 | 0.67358 | 3.8715 | 0.62435 | 1.2905
| 14.0333 | 0.67715 | 3.8716 | 0.62803 | 1.2905
| 14.15 | 0.67715 | 3.8717 | 0.62479 | 1.2906
| 14.1333 | 0.67358 | 3.872 | 0.62524 | 1.2907
| 14.1167 | 0.67358 | 3.8723 | 0.62568 | 1.2908
| 14.1 | 0.67715 | 3.8724 | 0.62612 | 1.2908
| 13.9833 | 0.67358 | 3.8724 | 0.62935 | 1.2908
| 14.0833 | 0.67358 | 3.8726 | 0.62656 | 1.2909
| 13.9667 | 0.67358 | 3.8727 | 0.6298 | 1.2909
| 13.95 | 0.67358 | 3.8728 | 0.63024 | 1.2909
| 14.0667 | 0.67358 | 3.873 | 0.627 | 1.291
| 14.05 | 0.67358 | 3.873 | 0.62745 | 1.291
| 13.9333 | 0.67358 | 3.8731 | 0.63068 | 1.291
| 14.0333 | 0.67358 | 3.8733 | 0.62789 | 1.2911
| 13.9 | 0.67358 | 3.8736 | 0.63156 | 1.2912
| 14.0167 | 0.67358 | 3.8737 | 0.62833 | 1.2912
| 14 | 0.67715 | 3.8738 | 0.62877 | 1.2913
| 13.8833 | 0.67358 | 3.874 | 0.63201 | 1.2913
| 13.9833 | 0.67715 | 3.8742 | 0.62921 | 1.2914
| 13.8667 | 0.67358 | 3.8743 | 0.63245 | 1.2914
| 13.9667 | 0.67358 | 3.8743 | 0.62966 | 1.2914
| 13.85 | 0.67715 | 3.8745 | 0.63289 | 1.2915
| 13.95 | 0.67715 | 3.8746 | 0.6301 | 1.2915
| 13.9333 | 0.67358 | 3.8748 | 0.63054 | 1.2916
| 13.8167 | 0.67001 | 3.875 | 0.63377 | 1.2917
| 13.9167 | 0.67358 | 3.8751 | 0.63098 | 1.2917
| 13.8 | 0.67715 | 3.8752 | 0.63421 | 1.2917
| 13.9 | 0.67358 | 3.8753 | 0.63142 | 1.2918
| 13.7833 | 0.67358 | 3.8755 | 0.63466 | 1.2918
| 13.8833 | 0.67715 | 3.8756 | 0.63187 | 1.2919
| 13.7667 | 0.67358 | 3.8757 | 0.6351 | 1.2919
| 13.8667 | 0.67358 | 3.876 | 0.63231 | 1.292
| 13.85 | 0.67001 | 3.8762 | 0.63275 | 1.2921
| 13.7333 | 0.67715 | 3.8763 | 0.63598 | 1.2921
| 13.8333 | 0.67358 | 3.8765 | 0.63319 | 1.2922
| 13.7167 | 0.67715 | 3.8765 | 0.63642 | 1.2922
| 13.8167 | 0.67358 | 3.8766 | 0.63363 | 1.2922
| 13.8 | 0.67715 | 3.8769 | 0.63408 | 1.2923
| 13.7 | 0.67358 | 3.8769 | 0.63686 | 1.2923
| 13.6833 | 0.67358 | 3.8772 | 0.63731 | 1.2924
| 13.7833 | 0.67358 | 3.8773 | 0.63452 | 1.2924
| 13.7667 | 0.67358 | 3.8774 | 0.63496 | 1.2925
| 13.75 | 0.67358 | 3.8778 | 0.6354 | 1.2926
| 13.7333 | 0.67715 | 3.8781 | 0.63584 | 1.2927
| 13.7167 | 0.67358 | 3.8783 | 0.63629 | 1.2928
| 13.6 | 0.67001 | 3.8785 | 0.63952 | 1.2928
| 13.7 | 0.67358 | 3.8786 | 0.63673 | 1.2929
| 13.6833 | 0.67358 | 3.8788 | 0.63717 | 1.2929
| 13.5667 | 0.67715 | 3.879 | 0.6404 | 1.293
| 13.6667 | 0.67358 | 3.8791 | 0.63761 | 1.293
| 13.65 | 0.67358 | 3.8794 | 0.63805 | 1.2931
| 13.6333 | 0.67358 | 3.8795 | 0.63849 | 1.2932
| 13.5333 | 0.67358 | 3.8797 | 0.64128 | 1.2932
| 13.6167 | 0.67358 | 3.8798 | 0.63894 | 1.2933
| 13.5167 | 0.67001 | 3.88 | 0.64172 | 1.2933
| 13.6 | 0.67358 | 3.8801 | 0.63938 | 1.2934
| 13.5 | 0.67358 | 3.8804 | 0.64217 | 1.2935
| 13.5833 | 0.67358 | 3.8806 | 0.63982 | 1.2935
| 13.4833 | 0.67358 | 3.8807 | 0.64261 | 1.2936
| 13.5667 | 0.67715 | 3.8808 | 0.64026 | 1.2936
| 13.55 | 0.67715 | 3.881 | 0.6407 | 1.2937
| 13.45 | 0.67358 | 3.8811 | 0.64349 | 1.2937
| 13.5333 | 0.67358 | 3.8813 | 0.64115 | 1.2938
| 13.4333 | 0.67715 | 3.8814 | 0.64393 | 1.2938
| 13.5167 | 0.67358 | 3.8817 | 0.64159 | 1.2939
| 13.4167 | 0.67358 | 3.8817 | 0.64437 | 1.2939
| 13.5 | 0.67715 | 3.882 | 0.64203 | 1.294
| 13.4833 | 0.67358 | 3.8821 | 0.64247 | 1.294
| 13.4667 | 0.67358 | 3.8824 | 0.64291 | 1.2941
| 13.3833 | 0.67358 | 3.8824 | 0.64526 | 1.2941
| 13.45 | 0.67715 | 3.8827 | 0.64336 | 1.2942
| 13.3667 | 0.67358 | 3.8828 | 0.6457 | 1.2943
| 13.4333 | 0.67715 | 3.883 | 0.6438 | 1.2943
| 13.4167 | 0.67358 | 3.8833 | 0.64424 | 1.2944
| 13.3167 | 0.67358 | 3.8836 | 0.64703 | 1.2945
| 13.4 | 0.67358 | 3.8836 | 0.64468 | 1.2945
| 13.3833 | 0.67715 | 3.8839 | 0.64512 | 1.2946
| 13.3667 | 0.67358 | 3.8842 | 0.64557 | 1.2947
| 13.35 | 0.67715 | 3.8844 | 0.64601 | 1.2948
| 13.2667 | 0.67715 | 3.8846 | 0.64835 | 1.2949
| 13.3333 | 0.67358 | 3.8849 | 0.64645 | 1.295
| 13.3167 | 0.67715 | 3.8849 | 0.64689 | 1.295
| 13.3 | 0.67358 | 3.8853 | 0.64733 | 1.2951
| 13.2833 | 0.67715 | 3.8856 | 0.64778 | 1.2952
| 13.2 | 0.67358 | 3.8858 | 0.65012 | 1.2953
| 13.2667 | 0.67358 | 3.886 | 0.64822 | 1.2953
| 13.1833 | 0.67358 | 3.8862 | 0.65056 | 1.2954
| 13.25 | 0.67358 | 3.8865 | 0.64866 | 1.2955
| 13.1667 | 0.67001 | 3.8866 | 0.651 | 1.2955
| 13.2333 | 0.67715 | 3.8866 | 0.6491 | 1.2955
| 13.15 | 0.67715 | 3.8868 | 0.65144 | 1.2956
| 13.2167 | 0.67715 | 3.8869 | 0.64954 | 1.2956
| 13.2 | 0.67358 | 3.8871 | 0.64999 | 1.2957
| 13.1333 | 0.67001 | 3.8872 | 0.65188 | 1.2957
| 13.1833 | 0.67358 | 3.8875 | 0.65043 | 1.2958
| 13.1667 | 0.67715 | 3.8877 | 0.65087 | 1.2959
| 13.1 | 0.67358 | 3.8879 | 0.65277 | 1.296
| 13.15 | 0.67715 | 3.8881 | 0.65131 | 1.296
| 13.0833 | 0.67358 | 3.8881 | 0.65321 | 1.296
| 13.1333 | 0.67358 | 3.8883 | 0.65175 | 1.2961
| 13.0667 | 0.67715 | 3.8885 | 0.65365 | 1.2962
| 13.1167 | 0.67715 | 3.8888 | 0.65219 | 1.2963
| 13.05 | 0.67358 | 3.8889 | 0.65409 | 1.2963
| 13.1 | 0.67715 | 3.8892 | 0.65264 | 1.2964
| 13.0833 | 0.67358 | 3.8894 | 0.65308 | 1.2965
| 13.0167 | 0.67358 | 3.8895 | 0.65498 | 1.2965
| 13.0667 | 0.67715 | 3.8897 | 0.65352 | 1.2966
| 13.05 | 0.67715 | 3.8898 | 0.65396 | 1.2966
| 13.0333 | 0.67358 | 3.8902 | 0.6544 | 1.2967
| 13.0167 | 0.67358 | 3.8905 | 0.65485 | 1.2968
| 12.9667 | 0.67358 | 3.8907 | 0.6563 | 1.2969
| 13 | 0.67358 | 3.8909 | 0.65529 | 1.297
| 12.95 | 0.67358 | 3.891 | 0.65674 | 1.297
| 12.9833 | 0.67358 | 3.8913 | 0.65573 | 1.2971
| 12.9333 | 0.67715 | 3.8914 | 0.65719 | 1.2971
| 12.9667 | 0.67715 | 3.8917 | 0.65617 | 1.2972
| 12.9 | 0.67715 | 3.8919 | 0.65807 | 1.2973
| 12.95 | 0.67715 | 3.892 | 0.65661 | 1.2973
| 12.8833 | 0.67358 | 3.8923 | 0.65851 | 1.2974
| 12.9333 | 0.67358 | 3.8923 | 0.65706 | 1.2974
| 12.9167 | 0.67358 | 3.8927 | 0.6575 | 1.2975
| 12.8667 | 0.67715 | 3.8928 | 0.65895 | 1.2976
| 12.9 | 0.67358 | 3.8929 | 0.65794 | 1.2976
| 12.85 | 0.67358 | 3.8932 | 0.65939 | 1.2977
| 12.8833 | 0.67358 | 3.8933 | 0.65838 | 1.2978
| 12.8333 | 0.67715 | 3.8934 | 0.65984 | 1.2978
| 12.8667 | 0.67715 | 3.8936 | 0.65882 | 1.2979
| 12.8167 | 0.67358 | 3.8938 | 0.66028 | 1.2979
| 12.85 | 0.67358 | 3.8939 | 0.65927 | 1.298
| 12.8333 | 0.67715 | 3.8942 | 0.65971 | 1.2981
| 12.8167 | 0.67715 | 3.8946 | 0.66015 | 1.2982
| 12.8 | 0.67358 | 3.8949 | 0.66059 | 1.2983
| 12.7667 | 0.67715 | 3.895 | 0.6616 | 1.2983
| 12.7833 | 0.67358 | 3.8953 | 0.66103 | 1.2984
| 12.7333 | 0.67715 | 3.8957 | 0.66249 | 1.2986
| 12.7667 | 0.67358 | 3.8958 | 0.66148 | 1.2986
| 12.75 | 0.67358 | 3.8959 | 0.66192 | 1.2986
| 12.7167 | 0.67715 | 3.896 | 0.66293 | 1.2987
| 12.7333 | 0.67358 | 3.8962 | 0.66236 | 1.2987
| 12.7 | 0.67715 | 3.8963 | 0.66337 | 1.2988
| 12.7167 | 0.67358 | 3.8966 | 0.6628 | 1.2989
| 12.6833 | 0.67715 | 3.8968 | 0.66381 | 1.2989
| 12.7 | 0.67358 | 3.8969 | 0.66324 | 1.299
| 12.6667 | 0.67715 | 3.8972 | 0.66425 | 1.2991
| 12.6833 | 0.67715 | 3.8973 | 0.66369 | 1.2991
| 12.6667 | 0.67715 | 3.8976 | 0.66413 | 1.2992
| 12.6333 | 0.67358 | 3.8979 | 0.66514 | 1.2993
| 12.65 | 0.67358 | 3.8981 | 0.66457 | 1.2994
| 12.6167 | 0.67715 | 3.8982 | 0.66558 | 1.2994
| 12.6333 | 0.67715 | 3.8984 | 0.66501 | 1.2995
| 12.6167 | 0.67715 | 3.8987 | 0.66545 | 1.2996
| 12.6 | 0.67358 | 3.899 | 0.66589 | 1.2997
| 12.5833 | 0.67358 | 3.8991 | 0.66646 | 1.2997
| 12.5833 | 0.67715 | 3.8994 | 0.66634 | 1.2998
| 12.5667 | 0.67715 | 3.8994 | 0.66691 | 1.2998
| 12.5667 | 0.67715 | 3.8998 | 0.66678 | 1.2999
| 12.55 | 0.67358 | 3.9001 | 0.66722 | 1.3
| 12.5333 | 0.67358 | 3.9002 | 0.66779 | 1.3001
| 12.5333 | 0.67358 | 3.9005 | 0.66766 | 1.3002
| 12.5167 | 0.67358 | 3.9007 | 0.66823 | 1.3002
| 12.5167 | 0.67715 | 3.9009 | 0.6681 | 1.3003
| 12.5 | 0.67715 | 3.901 | 0.66867 | 1.3003
| 12.5 | 0.67358 | 3.9012 | 0.66855 | 1.3004
| 12.4833 | 0.67715 | 3.9014 | 0.66911 | 1.3005
| 12.4833 | 0.67358 | 3.9016 | 0.66899 | 1.3005
| 12.4667 | 0.67715 | 3.9018 | 0.66943 | 1.3006
| 12.45 | 0.67358 | 3.9021 | 0.66987 | 1.3007
| 12.4333 | 0.67358 | 3.9024 | 0.67031 | 1.3008
| 12.4167 | 0.67715 | 3.9027 | 0.67088 | 1.3009
| 12.4167 | 0.67358 | 3.9029 | 0.67076 | 1.301
| 12.4 | 0.67358 | 3.9033 | 0.6712 | 1.3011
| 12.3833 | 0.67358 | 3.9036 | 0.67176 | 1.3012
| 12.3833 | 0.67358 | 3.9038 | 0.67164 | 1.3013
| 12.3667 | 0.67358 | 3.9039 | 0.67208 | 1.3013
| 12.3667 | 0.67715 | 3.904 | 0.67221 | 1.3013
| 12.35 | 0.67358 | 3.9045 | 0.67265 | 1.3015
| 12.35 | 0.67715 | 3.9045 | 0.67252 | 1.3015
| 12.3333 | 0.67715 | 3.9048 | 0.67297 | 1.3016
| 12.3167 | 0.67715 | 3.9052 | 0.67341 | 1.3017
| 12.3167 | 0.67358 | 3.9052 | 0.67353 | 1.3017
| 12.3 | 0.67715 | 3.9054 | 0.67385 | 1.3018
| 12.3 | 0.67358 | 3.9056 | 0.67397 | 1.3019
| 12.2833 | 0.67358 | 3.9059 | 0.67429 | 1.302
| 12.2667 | 0.67715 | 3.9061 | 0.67486 | 1.302
| 12.2667 | 0.67715 | 3.9062 | 0.67473 | 1.3021
| 12.25 | 0.67715 | 3.9064 | 0.67518 | 1.3021
| 12.25 | 0.67358 | 3.9066 | 0.6753 | 1.3022
| 12.2333 | 0.67358 | 3.9068 | 0.67562 | 1.3023
| 12.2333 | 0.67358 | 3.9069 | 0.67574 | 1.3023
| 12.2167 | 0.67715 | 3.9072 | 0.67606 | 1.3024
| 12.2167 | 0.67715 | 3.9074 | 0.67618 | 1.3025
| 12.2 | 0.67358 | 3.9076 | 0.6765 | 1.3025
| 12.2 | 0.67715 | 3.9079 | 0.67662 | 1.3026
| 12.1833 | 0.67715 | 3.908 | 0.67694 | 1.3027
| 12.1833 | 0.67358 | 3.9083 | 0.67707 | 1.3028
| 12.1667 | 0.67358 | 3.9085 | 0.67739 | 1.3028
| 12.1667 | 0.67358 | 3.9086 | 0.67751 | 1.3029
| 12.15 | 0.67358 | 3.9087 | 0.67783 | 1.3029
| 12.15 | 0.67715 | 3.9089 | 0.67795 | 1.303
| 12.1333 | 0.67715 | 3.9091 | 0.67827 | 1.303
| 12.1333 | 0.67358 | 3.9092 | 0.67839 | 1.3031
| 12.1167 | 0.67358 | 3.9095 | 0.67871 | 1.3032
| 12.1167 | 0.67358 | 3.9096 | 0.67883 | 1.3032
| 12.1 | 0.67715 | 3.91 | 0.67915 | 1.3033
| 12.0833 | 0.67358 | 3.9103 | 0.6796 | 1.3034
| 12.0667 | 0.67715 | 3.9107 | 0.68004 | 1.3036
| 12.0667 | 0.67715 | 3.9108 | 0.68016 | 1.3036
| 12.05 | 0.67358 | 3.911 | 0.68048 | 1.3037
| 12.05 | 0.67715 | 3.9112 | 0.6806 | 1.3037
| 12.0333 | 0.67715 | 3.9113 | 0.68092 | 1.3038
| 12.0167 | 0.67715 | 3.9117 | 0.68136 | 1.3039
| 12.0167 | 0.67358 | 3.912 | 0.68148 | 1.304
| 12 | 0.67715 | 3.9122 | 0.6818 | 1.3041
| 12 | 0.67358 | 3.9123 | 0.68193 | 1.3041
| 11.9833 | 0.67358 | 3.9126 | 0.68225 | 1.3042
| 11.9833 | 0.67715 | 3.9127 | 0.68237 | 1.3042
| 11.9667 | 0.67358 | 3.913 | 0.68269 | 1.3043
| 11.95 | 0.67358 | 3.9132 | 0.68325 | 1.3044
| 11.95 | 0.67358 | 3.9133 | 0.68313 | 1.3044
| 11.9333 | 0.67358 | 3.9136 | 0.68357 | 1.3045
| 11.9333 | 0.67358 | 3.9137 | 0.68369 | 1.3046
| 11.9167 | 0.67358 | 3.9141 | 0.68401 | 1.3047
| 11.9 | 0.67715 | 3.9143 | 0.68446 | 1.3048
| 11.9 | 0.67715 | 3.9145 | 0.68458 | 1.3048
| 11.8833 | 0.67358 | 3.9147 | 0.6849 | 1.3049
| 11.8833 | 0.67358 | 3.9149 | 0.68502 | 1.305
| 11.8667 | 0.67358 | 3.9151 | 0.68534 | 1.305
| 11.8667 | 0.67358 | 3.9152 | 0.68546 | 1.3051
| 11.85 | 0.67715 | 3.9155 | 0.68578 | 1.3052
| 11.85 | 0.67358 | 3.9155 | 0.6859 | 1.3052
| 11.8333 | 0.67715 | 3.9159 | 0.68622 | 1.3053
| 11.8333 | 0.67715 | 3.916 | 0.68634 | 1.3053
| 11.8167 | 0.67358 | 3.9162 | 0.68667 | 1.3054
| 11.8 | 0.67715 | 3.9165 | 0.68711 | 1.3055
| 11.7833 | 0.67358 | 3.917 | 0.68767 | 1.3057
| 11.7833 | 0.67715 | 3.9171 | 0.68755 | 1.3057
| 11.7667 | 0.67715 | 3.9172 | 0.68799 | 1.3057
| 11.7667 | 0.67358 | 3.9173 | 0.68811 | 1.3058
| 11.75 | 0.67358 | 3.9177 | 0.68843 | 1.3059
| 11.75 | 0.67358 | 3.9178 | 0.68855 | 1.3059
| 11.7333 | 0.67715 | 3.9179 | 0.68888 | 1.306
| 11.7333 | 0.67715 | 3.9181 | 0.68899 | 1.306
| 11.7167 | 0.67715 | 3.9182 | 0.68932 | 1.3061
| 11.7 | 0.67715 | 3.9185 | 0.68976 | 1.3062
| 11.7 | 0.67715 | 3.9188 | 0.68988 | 1.3063
| 11.6833 | 0.67358 | 3.919 | 0.6902 | 1.3063
| 11.6833 | 0.67715 | 3.9192 | 0.69032 | 1.3064
| 11.6667 | 0.67715 | 3.9195 | 0.69064 | 1.3065
| 11.65 | 0.67358 | 3.9197 | 0.69109 | 1.3066
| 11.65 | 0.67358 | 3.9199 | 0.6912 | 1.3066
| 11.6333 | 0.67715 | 3.9201 | 0.69164 | 1.3067
| 11.6333 | 0.67358 | 3.9202 | 0.69153 | 1.3067
| 11.6167 | 0.67358 | 3.9205 | 0.69197 | 1.3068
| 11.6 | 0.67358 | 3.9207 | 0.69241 | 1.3069
| 11.6 | 0.67715 | 3.921 | 0.69253 | 1.307
| 11.5833 | 0.67358 | 3.921 | 0.69285 | 1.307
| 11.5833 | 0.67715 | 3.9213 | 0.69297 | 1.3071
| 11.5667 | 0.67715 | 3.9213 | 0.6933 | 1.3071
| 11.5667 | 0.67715 | 3.9216 | 0.69341 | 1.3072
| 11.55 | 0.67001 | 3.9219 | 0.69374 | 1.3073
| 11.55 | 0.67358 | 3.9219 | 0.69385 | 1.3073
| 11.5333 | 0.67358 | 3.9221 | 0.69418 | 1.3074
| 11.5333 | 0.67358 | 3.9223 | 0.69429 | 1.3074
| 11.5167 | 0.67715 | 3.9226 | 0.69462 | 1.3075
| 11.5 | 0.67715 | 3.9229 | 0.69506 | 1.3076
| 11.5 | 0.67358 | 3.923 | 0.69518 | 1.3077
| 11.4833 | 0.67358 | 3.9232 | 0.6955 | 1.3077
| 11.4667 | 0.67358 | 3.9235 | 0.69595 | 1.3078
| 11.4667 | 0.67715 | 3.9236 | 0.69606 | 1.3079
| 11.45 | 0.67358 | 3.9238 | 0.69639 | 1.3079
| 11.45 | 0.67358 | 3.9241 | 0.6965 | 1.308
| 11.4333 | 0.67715 | 3.9242 | 0.69683 | 1.3081
| 11.4167 | 0.67715 | 3.9245 | 0.69727 | 1.3082
| 11.4167 | 0.67358 | 3.9249 | 0.69739 | 1.3083
| 11.4 | 0.67358 | 3.9249 | 0.69771 | 1.3083
| 11.4 | 0.67358 | 3.9251 | 0.69783 | 1.3084
| 11.3833 | 0.67715 | 3.9253 | 0.69816 | 1.3084
| 11.3833 | 0.67358 | 3.9255 | 0.69827 | 1.3085
| 11.3667 | 0.67715 | 3.9255 | 0.6986 | 1.3085
| 11.3667 | 0.67715 | 3.9257 | 0.69871 | 1.3086
| 11.35 | 0.67358 | 3.926 | 0.69904 | 1.3087
| 11.35 | 0.67358 | 3.9261 | 0.69915 | 1.3087
| 11.3333 | 0.67715 | 3.9264 | 0.69948 | 1.3088
| 11.3333 | 0.67715 | 3.9264 | 0.6996 | 1.3088
| 11.3167 | 0.67358 | 3.9265 | 0.69992 | 1.3088
| 11.3167 | 0.67358 | 3.9267 | 0.70004 | 1.3089
| 11.3 | 0.67358 | 3.9268 | 0.70037 | 1.3089
| 11.3 | 0.67358 | 3.9271 | 0.70048 | 1.309
| 11.2833 | 0.67358 | 3.9272 | 0.70081 | 1.3091
| 11.2833 | 0.67715 | 3.9273 | 0.70092 | 1.3091
| 11.2667 | 0.67358 | 3.9275 | 0.70125 | 1.3092
| 11.2667 | 0.67715 | 3.9277 | 0.70136 | 1.3092
| 11.25 | 0.67715 | 3.928 | 0.70169 | 1.3093
| 11.25 | 0.67358 | 3.9281 | 0.7018 | 1.3094
| 11.2333 | 0.67358 | 3.9283 | 0.70213 | 1.3094
| 11.2333 | 0.67715 | 3.9284 | 0.70225 | 1.3095
| 11.2167 | 0.67715 | 3.9287 | 0.70258 | 1.3096
| 11.2167 | 0.67358 | 3.9288 | 0.70269 | 1.3096
| 11.2 | 0.67358 | 3.929 | 0.70302 | 1.3097
| 11.2 | 0.67358 | 3.9291 | 0.70313 | 1.3097
| 11.1833 | 0.67715 | 3.9293 | 0.70346 | 1.3098
| 11.1833 | 0.67358 | 3.9295 | 0.70357 | 1.3098
| 11.1667 | 0.67358 | 3.9297 | 0.7039 | 1.3099
| 11.1667 | 0.67001 | 3.9298 | 0.70401 | 1.3099
| 11.15 | 0.67358 | 3.9299 | 0.70434 | 1.31
| 11.15 | 0.67358 | 3.93 | 0.70446 | 1.31
| 11.1333 | 0.67715 | 3.9302 | 0.70479 | 1.3101
| 11.1333 | 0.67358 | 3.9304 | 0.7049 | 1.3101
| 11.1167 | 0.67358 | 3.9305 | 0.70523 | 1.3102
| 11.1167 | 0.67001 | 3.9307 | 0.70534 | 1.3102
| 11.1 | 0.67358 | 3.931 | 0.70567 | 1.3103
| 11.1 | 0.67358 | 3.9311 | 0.70578 | 1.3104
| 11.0833 | 0.67715 | 3.9313 | 0.70611 | 1.3104
| 11.0833 | 0.67715 | 3.9315 | 0.70622 | 1.3105
| 11.0667 | 0.67715 | 3.9316 | 0.70655 | 1.3105
| 11.0667 | 0.67358 | 3.9317 | 0.70666 | 1.3106
| 11.05 | 0.67358 | 3.932 | 0.707 | 1.3107
| 11.05 | 0.67715 | 3.9322 | 0.70711 | 1.3107
| 11.0333 | 0.67358 | 3.9322 | 0.70744 | 1.3107
| 11.0333 | 0.67001 | 3.9326 | 0.70755 | 1.3109
| 11.0167 | 0.67358 | 3.9326 | 0.70788 | 1.3109
| 11.0167 | 0.67358 | 3.9328 | 0.70799 | 1.3109
| 11 | 0.67358 | 3.9329 | 0.70832 | 1.311
| 11 | 0.68072 | 3.933 | 0.70843 | 1.311
| 10.9833 | 0.67358 | 3.9332 | 0.70876 | 1.3111
| 10.9667 | 0.67358 | 3.9336 | 0.7092 | 1.3112
| 10.9667 | 0.67001 | 3.9337 | 0.70931 | 1.3112
| 10.95 | 0.67358 | 3.934 | 0.70965 | 1.3113
| 10.95 | 0.67358 | 3.9341 | 0.70976 | 1.3114
| 10.9333 | 0.67715 | 3.9342 | 0.71009 | 1.3114
| 10.9333 | 0.67715 | 3.9345 | 0.7102 | 1.3115
| 10.9167 | 0.67358 | 3.9345 | 0.71053 | 1.3115
| 10.9167 | 0.67358 | 3.9347 | 0.71064 | 1.3116
| 10.9 | 0.67358 | 3.935 | 0.71097 | 1.3117
| 10.9 | 0.67358 | 3.9351 | 0.71108 | 1.3117
| 10.8833 | 0.67715 | 3.9352 | 0.71141 | 1.3117
| 10.8833 | 0.67358 | 3.9353 | 0.71152 | 1.3118
| 10.8667 | 0.67358 | 3.9355 | 0.71186 | 1.3118
| 10.8667 | 0.67358 | 3.9357 | 0.71197 | 1.3119
| 10.85 | 0.67358 | 3.9358 | 0.7123 | 1.3119
| 10.85 | 0.67715 | 3.9358 | 0.71241 | 1.3119
| 10.8333 | 0.67358 | 3.9361 | 0.71274 | 1.312
| 10.8167 | 0.67715 | 3.9365 | 0.71318 | 1.3122
| 10.8167 | 0.67715 | 3.9367 | 0.71329 | 1.3122
| 10.8 | 0.67358 | 3.9368 | 0.71362 | 1.3123
| 10.8 | 0.67715 | 3.9369 | 0.71373 | 1.3123
| 10.7833 | 0.67358 | 3.9371 | 0.71407 | 1.3124
| 10.7833 | 0.67358 | 3.9373 | 0.71417 | 1.3124
| 10.7667 | 0.67715 | 3.9374 | 0.71462 | 1.3125
| 10.7667 | 0.67358 | 3.9376 | 0.71451 | 1.3125
| 10.75 | 0.67358 | 3.9378 | 0.71495 | 1.3126
| 10.7333 | 0.67358 | 3.9381 | 0.71539 | 1.3127
| 10.7167 | 0.67715 | 3.9384 | 0.71583 | 1.3128
| 10.7167 | 0.67715 | 3.9385 | 0.71594 | 1.3128
| 10.7 | 0.67358 | 3.9387 | 0.71638 | 1.3129
| 10.7 | 0.67358 | 3.9388 | 0.71628 | 1.3129
| 10.6833 | 0.67715 | 3.939 | 0.71683 | 1.313
| 10.6833 | 0.67715 | 3.9391 | 0.71672 | 1.313
| 10.6667 | 0.67358 | 3.9394 | 0.71727 | 1.3131
| 10.6667 | 0.67358 | 3.9394 | 0.71716 | 1.3131
| 10.65 | 0.67715 | 3.9396 | 0.71771 | 1.3132
| 10.65 | 0.67358 | 3.9396 | 0.7176 | 1.3132
| 10.6333 | 0.67358 | 3.9398 | 0.71815 | 1.3133
| 10.6333 | 0.67715 | 3.9401 | 0.71804 | 1.3134
| 10.6167 | 0.67715 | 3.9403 | 0.71859 | 1.3134
| 10.6167 | 0.67358 | 3.9404 | 0.71849 | 1.3135
| 10.6 | 0.67358 | 3.9406 | 0.71903 | 1.3135
| 10.6 | 0.67358 | 3.9408 | 0.71893 | 1.3136
| 10.5833 | 0.67715 | 3.9409 | 0.71948 | 1.3136
| 10.5833 | 0.67358 | 3.941 | 0.71937 | 1.3137
| 10.5667 | 0.67715 | 3.9411 | 0.71992 | 1.3137
| 10.5667 | 0.67358 | 3.9413 | 0.71981 | 1.3137
| 10.55 | 0.67358 | 3.9415 | 0.72036 | 1.3138
| 10.55 | 0.67358 | 3.9416 | 0.72025 | 1.3139
| 10.5333 | 0.67358 | 3.9417 | 0.7208 | 1.3139
| 10.5333 | 0.67358 | 3.9419 | 0.7207 | 1.314
| 10.5167 | 0.67715 | 3.942 | 0.72124 | 1.314
| 10.5167 | 0.67358 | 3.9421 | 0.72114 | 1.314
| 10.5 | 0.67358 | 3.9422 | 0.72168 | 1.3141
| 10.5 | 0.67358 | 3.9426 | 0.72158 | 1.3142
| 10.4833 | 0.67358 | 3.9427 | 0.72213 | 1.3142
| 10.4833 | 0.67358 | 3.9428 | 0.72202 | 1.3143
| 10.4667 | 0.67358 | 3.943 | 0.72257 | 1.3143
| 10.4667 | 0.67715 | 3.9433 | 0.72246 | 1.3144
| 10.45 | 0.67358 | 3.9435 | 0.72301 | 1.3145
| 10.45 | 0.67715 | 3.9435 | 0.7229 | 1.3145
| 10.4333 | 0.67715 | 3.9437 | 0.72345 | 1.3146
| 10.4333 | 0.67358 | 3.9438 | 0.72335 | 1.3146
| 10.4167 | 0.67358 | 3.9441 | 0.72379 | 1.3147
| 10.4 | 0.67715 | 3.9442 | 0.72434 | 1.3147
| 10.4 | 0.67715 | 3.9444 | 0.72423 | 1.3148
| 10.3833 | 0.67358 | 3.9445 | 0.72478 | 1.3148
| 10.3833 | 0.67715 | 3.9447 | 0.72467 | 1.3149
| 10.3667 | 0.67358 | 3.9448 | 0.72522 | 1.3149
| 10.3667 | 0.67358 | 3.945 | 0.72511 | 1.315
| 10.35 | 0.67715 | 3.9453 | 0.72556 | 1.3151
| 10.3333 | 0.67715 | 3.9457 | 0.726 | 1.3153
| 10.3167 | 0.67358 | 3.946 | 0.72644 | 1.3153
| 10.3 | 0.67358 | 3.9463 | 0.72688 | 1.3154
| 10.2833 | 0.67715 | 3.9464 | 0.72743 | 1.3155
| 10.2833 | 0.67715 | 3.9466 | 0.72732 | 1.3155
| 10.25 | 0.67715 | 3.9468 | 0.72831 | 1.3156
| 10.2667 | 0.67001 | 3.9469 | 0.72777 | 1.3156
| 10.2333 | 0.67001 | 3.9471 | 0.72875 | 1.3157
| 10.25 | 0.67715 | 3.9472 | 0.72821 | 1.3157
| 10.2167 | 0.67358 | 3.9474 | 0.72919 | 1.3158
| 10.2333 | 0.67358 | 3.9474 | 0.72865 | 1.3158
| 10.2 | 0.67001 | 3.9477 | 0.72964 | 1.3159
| 10.2167 | 0.67358 | 3.9478 | 0.72909 | 1.3159
| 10.1833 | 0.67358 | 3.948 | 0.73008 | 1.316
| 10.2 | 0.67358 | 3.9481 | 0.72953 | 1.316
| 10.1667 | 0.67715 | 3.9483 | 0.73052 | 1.3161
| 10.1833 | 0.67358 | 3.9485 | 0.72998 | 1.3162
| 10.15 | 0.67001 | 3.9487 | 0.73096 | 1.3162
| 10.1667 | 0.67358 | 3.9488 | 0.73042 | 1.3163
| 10.15 | 0.67358 | 3.949 | 0.73086 | 1.3163
| 10.1167 | 0.67358 | 3.9493 | 0.73185 | 1.3164
| 10.1333 | 0.67358 | 3.9494 | 0.7313 | 1.3165
| 10.1 | 0.67358 | 3.9497 | 0.73229 | 1.3166
| 10.1167 | 0.67358 | 3.9497 | 0.73174 | 1.3166
| 10.0833 | 0.67715 | 3.9499 | 0.73273 | 1.3166
| 10.1 | 0.67358 | 3.9499 | 0.73219 | 1.3166
| 10.0667 | 0.67358 | 3.9501 | 0.73317 | 1.3167
| 10.0833 | 0.67715 | 3.9503 | 0.73263 | 1.3168
| 10.05 | 0.67358 | 3.9504 | 0.73361 | 1.3168
| 10.0667 | 0.67715 | 3.9505 | 0.73307 | 1.3168
| 10.0333 | 0.67358 | 3.9507 | 0.73405 | 1.3169
| 10.05 | 0.67715 | 3.9509 | 0.73351 | 1.317
| 10.0333 | 0.67715 | 3.9512 | 0.73395 | 1.3171
| 10 | 0.67715 | 3.9514 | 0.73494 | 1.3171
| 10.0167 | 0.67001 | 3.9516 | 0.7344 | 1.3172
| 9.9833 | 0.67358 | 3.9517 | 0.73538 | 1.3172
| 10 | 0.67715 | 3.9519 | 0.73484 | 1.3173
| 9.9667 | 0.67358 | 3.952 | 0.73582 | 1.3173
| 9.9833 | 0.67358 | 3.9522 | 0.73528 | 1.3174
| 9.9667 | 0.67358 | 3.9524 | 0.73572 | 1.3175
| 9.9333 | 0.67715 | 3.9525 | 0.7367 | 1.3175
| 9.95 | 0.67358 | 3.9526 | 0.73616 | 1.3175
| 9.9333 | 0.67715 | 3.9529 | 0.7366 | 1.3176
| 9.9167 | 0.67715 | 3.9531 | 0.73705 | 1.3177
| 9.9 | 0.67715 | 3.9532 | 0.73759 | 1.3177
| 9.8833 | 0.67358 | 3.9534 | 0.73803 | 1.3178
| 9.9 | 0.67358 | 3.9536 | 0.73749 | 1.3179
| 9.8833 | 0.67715 | 3.9538 | 0.73793 | 1.3179
| 9.85 | 0.67358 | 3.9541 | 0.73891 | 1.318
| 9.8667 | 0.67715 | 3.9542 | 0.73837 | 1.3181
| 9.8333 | 0.67715 | 3.9544 | 0.73936 | 1.3181
| 9.85 | 0.67358 | 3.9545 | 0.73881 | 1.3182
| 9.8167 | 0.67358 | 3.9546 | 0.7398 | 1.3182
| 9.8333 | 0.67358 | 3.9549 | 0.73926 | 1.3183
| 9.8 | 0.67715 | 3.955 | 0.74024 | 1.3183
| 9.8167 | 0.67715 | 3.9551 | 0.7397 | 1.3184
| 9.7833 | 0.67715 | 3.9552 | 0.74068 | 1.3184
| 9.8 | 0.67358 | 3.9554 | 0.74014 | 1.3185
| 9.7667 | 0.67358 | 3.9555 | 0.74112 | 1.3185
| 9.7833 | 0.67358 | 3.9556 | 0.74058 | 1.3185
| 9.75 | 0.67358 | 3.9558 | 0.74156 | 1.3186
| 9.7667 | 0.67715 | 3.956 | 0.74102 | 1.3187
| 9.75 | 0.67715 | 3.9562 | 0.74147 | 1.3187
| 9.7333 | 0.67715 | 3.9563 | 0.74201 | 1.3188
| 9.7167 | 0.67358 | 3.9564 | 0.74245 | 1.3188
| 9.7333 | 0.67715 | 3.9567 | 0.74191 | 1.3189
| 9.7 | 0.67358 | 3.9567 | 0.74289 | 1.3189
| 9.7167 | 0.67715 | 3.9569 | 0.74235 | 1.319
| 9.6833 | 0.67358 | 3.9571 | 0.74333 | 1.319
| 9.7 | 0.67715 | 3.9572 | 0.74279 | 1.3191
| 9.6667 | 0.67715 | 3.9573 | 0.74377 | 1.3191
| 9.6833 | 0.67358 | 3.9575 | 0.74323 | 1.3192
| 9.6333 | 0.67715 | 3.9579 | 0.74466 | 1.3193
| 9.6667 | 0.67358 | 3.958 | 0.74368 | 1.3193
| 9.65 | 0.67358 | 3.958 | 0.74412 | 1.3193
| 9.6167 | 0.67715 | 3.9581 | 0.7451 | 1.3194
| 9.6333 | 0.67715 | 3.9583 | 0.74456 | 1.3194
| 9.6 | 0.67358 | 3.9584 | 0.74554 | 1.3195
| 9.6167 | 0.67358 | 3.9587 | 0.745 | 1.3196
| 9.5833 | 0.67715 | 3.9587 | 0.74598 | 1.3196
| 9.6 | 0.67715 | 3.959 | 0.74544 | 1.3197
| 9.5667 | 0.67358 | 3.9592 | 0.74642 | 1.3197
| 9.5833 | 0.67715 | 3.9593 | 0.74589 | 1.3198
| 9.55 | 0.67715 | 3.9595 | 0.74687 | 1.3198
| 9.5667 | 0.67358 | 3.9596 | 0.74633 | 1.3199
| 9.5333 | 0.67715 | 3.9597 | 0.74731 | 1.3199
| 9.5167 | 0.67715 | 3.96 | 0.74775 | 1.32
| 9.55 | 0.67715 | 3.96 | 0.74677 | 1.32
| 9.5 | 0.67358 | 3.9602 | 0.74819 | 1.3201
| 9.5333 | 0.67358 | 3.9603 | 0.74721 | 1.3201
| 9.5167 | 0.67358 | 3.9605 | 0.74765 | 1.3202
| 9.5 | 0.67715 | 3.9607 | 0.7481 | 1.3202
| 9.4667 | 0.67358 | 3.9608 | 0.74907 | 1.3203
| 9.4833 | 0.67358 | 3.9609 | 0.74854 | 1.3203
| 9.45 | 0.67358 | 3.9611 | 0.74952 | 1.3204
| 9.4667 | 0.67358 | 3.9613 | 0.74898 | 1.3204
| 9.4333 | 0.67715 | 3.9615 | 0.74996 | 1.3205
| 9.45 | 0.67358 | 3.9616 | 0.74942 | 1.3205
| 9.4333 | 0.67358 | 3.962 | 0.74986 | 1.3207
| 9.4 | 0.67358 | 3.9621 | 0.75084 | 1.3207
| 9.4167 | 0.67715 | 3.9623 | 0.75031 | 1.3208
| 9.4 | 0.67358 | 3.9626 | 0.75075 | 1.3209
| 9.3833 | 0.67715 | 3.9629 | 0.75119 | 1.321
| 9.3333 | 0.67358 | 3.9631 | 0.75261 | 1.321
| 9.3667 | 0.67358 | 3.9632 | 0.75163 | 1.3211
| 9.3167 | 0.67358 | 3.9634 | 0.75305 | 1.3211
| 9.35 | 0.67358 | 3.9634 | 0.75207 | 1.3212
| 9.3333 | 0.67358 | 3.9638 | 0.75251 | 1.3213
| 9.3167 | 0.67715 | 3.964 | 0.75296 | 1.3213
| 9.2667 | 0.67358 | 3.9643 | 0.75438 | 1.3214
| 9.3 | 0.67715 | 3.9644 | 0.7534 | 1.3215
| 9.2833 | 0.67358 | 3.9647 | 0.75384 | 1.3216
| 9.25 | 0.67715 | 3.9647 | 0.75482 | 1.3216
| 9.2333 | 0.67715 | 3.9649 | 0.75526 | 1.3216
| 9.2667 | 0.67358 | 3.965 | 0.75428 | 1.3217
| 9.2167 | 0.67715 | 3.9651 | 0.7557 | 1.3217
| 9.25 | 0.67358 | 3.9653 | 0.75472 | 1.3218
| 9.2 | 0.67358 | 3.9654 | 0.75614 | 1.3218
| 9.2333 | 0.67715 | 3.9655 | 0.75517 | 1.3218
| 9.2167 | 0.67715 | 3.9658 | 0.75561 | 1.3219
| 9.1667 | 0.67715 | 3.966 | 0.75703 | 1.322
| 9.2 | 0.67358 | 3.9661 | 0.75605 | 1.322
| 9.15 | 0.67715 | 3.9663 | 0.75747 | 1.3221
| 9.1833 | 0.67358 | 3.9664 | 0.75649 | 1.3221
| 9.1667 | 0.67358 | 3.9667 | 0.75693 | 1.3222
| 9.15 | 0.67358 | 3.967 | 0.75738 | 1.3223
| 9.1167 | 0.67715 | 3.967 | 0.75835 | 1.3223
| 9.1333 | 0.67358 | 3.9673 | 0.75782 | 1.3224
| 9.1 | 0.67715 | 3.9674 | 0.75879 | 1.3225
| 9.1167 | 0.67358 | 3.9676 | 0.75826 | 1.3225
| 9.1 | 0.67715 | 3.9679 | 0.7587 | 1.3226
| 9.0667 | 0.67715 | 3.968 | 0.75968 | 1.3227
| 9.05 | 0.67358 | 3.9681 | 0.76012 | 1.3227
| 9.0833 | 0.67715 | 3.9682 | 0.75914 | 1.3227
| 9.0333 | 0.67715 | 3.9684 | 0.76056 | 1.3228
| 9.0667 | 0.67358 | 3.9685 | 0.75959 | 1.3228
| 9.0167 | 0.67715 | 3.9687 | 0.761 | 1.3229
| 9.05 | 0.67358 | 3.9688 | 0.76003 | 1.3229
| 9.0333 | 0.67715 | 3.969 | 0.76047 | 1.323
| 8.9833 | 0.67358 | 3.9693 | 0.76189 | 1.3231
| 9.0167 | 0.67715 | 3.9694 | 0.76091 | 1.3231
| 8.9667 | 0.67358 | 3.9696 | 0.76233 | 1.3232
| 9 | 0.67358 | 3.9697 | 0.76135 | 1.3232
| 8.9833 | 0.67715 | 3.97 | 0.7618 | 1.3233
| 8.9333 | 0.67358 | 3.9701 | 0.76321 | 1.3234
| 8.9667 | 0.67358 | 3.9703 | 0.76224 | 1.3234
| 8.9167 | 0.67715 | 3.9705 | 0.76365 | 1.3235
| 8.95 | 0.67715 | 3.9707 | 0.76268 | 1.3236
| 8.9333 | 0.67358 | 3.9709 | 0.76312 | 1.3236
| 8.9167 | 0.67358 | 3.9712 | 0.76356 | 1.3237
| 8.8667 | 0.67715 | 3.9713 | 0.76498 | 1.3238
| 8.9 | 0.67358 | 3.9715 | 0.76401 | 1.3238
| 8.8833 | 0.67715 | 3.9716 | 0.76445 | 1.3239
| 8.8333 | 0.67358 | 3.9719 | 0.76586 | 1.324
| 8.8667 | 0.67358 | 3.9721 | 0.76489 | 1.324
| 8.8167 | 0.67358 | 3.9723 | 0.7663 | 1.3241
| 8.85 | 0.67358 | 3.9725 | 0.76533 | 1.3242
| 8.8333 | 0.67715 | 3.9726 | 0.76577 | 1.3242
| 8.7833 | 0.67715 | 3.9729 | 0.76719 | 1.3243
| 8.8167 | 0.67358 | 3.9731 | 0.76621 | 1.3244
| 8.7667 | 0.67715 | 3.9732 | 0.76763 | 1.3244
| 8.8 | 0.67358 | 3.9733 | 0.76666 | 1.3244
| 8.75 | 0.67358 | 3.9734 | 0.76807 | 1.3245
| 8.7833 | 0.67715 | 3.9737 | 0.7671 | 1.3246
| 8.7667 | 0.67358 | 3.974 | 0.76754 | 1.3247
| 8.75 | 0.67001 | 3.9741 | 0.76798 | 1.3247
| 8.7333 | 0.67715 | 3.9743 | 0.76842 | 1.3248
| 8.6833 | 0.67358 | 3.9745 | 0.76984 | 1.3248
| 8.7167 | 0.67358 | 3.9748 | 0.76887 | 1.3249
| 8.7 | 0.67715 | 3.9749 | 0.76931 | 1.325
| 8.65 | 0.67715 | 3.9752 | 0.77072 | 1.3251
| 8.6333 | 0.67358 | 3.9754 | 0.77116 | 1.3251
| 8.6833 | 0.67715 | 3.9755 | 0.76975 | 1.3252
| 8.6167 | 0.67715 | 3.9757 | 0.7716 | 1.3252
| 8.6667 | 0.67715 | 3.9757 | 0.77019 | 1.3252
| 8.65 | 0.67358 | 3.976 | 0.77063 | 1.3253
| 8.6 | 0.67358 | 3.9761 | 0.77205 | 1.3254
| 8.5833 | 0.67715 | 3.9764 | 0.77249 | 1.3255
| 8.6333 | 0.67715 | 3.9764 | 0.77108 | 1.3255
| 8.6167 | 0.67358 | 3.9766 | 0.77152 | 1.3255
| 8.5667 | 0.67358 | 3.9766 | 0.77293 | 1.3256
| 8.6 | 0.67358 | 3.9767 | 0.77196 | 1.3256
| 8.55 | 0.67358 | 3.9769 | 0.77337 | 1.3256
| 8.5833 | 0.67715 | 3.977 | 0.7724 | 1.3257
| 8.5333 | 0.67715 | 3.9773 | 0.77381 | 1.3258
| 8.5667 | 0.67358 | 3.9774 | 0.77284 | 1.3258
| 8.55 | 0.67715 | 3.9778 | 0.77329 | 1.3259
| 8.5 | 0.67358 | 3.9779 | 0.7747 | 1.326
| 8.4833 | 0.67358 | 3.9782 | 0.77514 | 1.3261
| 8.5333 | 0.67715 | 3.9783 | 0.77373 | 1.3261
| 8.4667 | 0.67358 | 3.9784 | 0.77558 | 1.3261
| 8.5167 | 0.67358 | 3.9785 | 0.77417 | 1.3262
| 8.5 | 0.67358 | 3.9787 | 0.77461 | 1.3262
| 8.4833 | 0.68072 | 3.9789 | 0.77505 | 1.3263
| 8.4333 | 0.67358 | 3.979 | 0.77646 | 1.3263
| 8.4667 | 0.67715 | 3.9792 | 0.7755 | 1.3264
| 8.45 | 0.67358 | 3.9796 | 0.77594 | 1.3265
| 8.4333 | 0.67715 | 3.9799 | 0.77638 | 1.3266
| 8.3833 | 0.67715 | 3.9799 | 0.77779 | 1.3266
| 8.4167 | 0.67715 | 3.9802 | 0.77682 | 1.3267
| 8.3667 | 0.67358 | 3.9803 | 0.77823 | 1.3268
| 8.4 | 0.67715 | 3.9805 | 0.77726 | 1.3268
| 8.3833 | 0.68072 | 3.9809 | 0.77771 | 1.327
| 8.3167 | 0.67358 | 3.9811 | 0.77956 | 1.327
| 8.3667 | 0.67715 | 3.9812 | 0.77815 | 1.3271
| 8.3 | 0.67715 | 3.9814 | 0.78 | 1.3271
| 8.35 | 0.67358 | 3.9815 | 0.77859 | 1.3272
| 8.2833 | 0.67001 | 3.9816 | 0.78044 | 1.3272
| 8.3333 | 0.67715 | 3.9817 | 0.77903 | 1.3272
| 8.2667 | 0.67358 | 3.982 | 0.78088 | 1.3273
| 8.3167 | 0.67358 | 3.9821 | 0.77947 | 1.3274
| 8.25 | 0.67715 | 3.9821 | 0.78132 | 1.3274
| 8.3 | 0.67358 | 3.9823 | 0.77991 | 1.3274
| 8.2833 | 0.67358 | 3.9826 | 0.78036 | 1.3275
| 8.2333 | 0.67358 | 3.9827 | 0.78177 | 1.3276
| 8.2667 | 0.67358 | 3.9829 | 0.7808 | 1.3276
| 8.2167 | 0.67358 | 3.983 | 0.78221 | 1.3277
| 8.25 | 0.67358 | 3.9832 | 0.78124 | 1.3277
| 8.1833 | 0.67715 | 3.9835 | 0.78309 | 1.3278
| 8.2333 | 0.67358 | 3.9836 | 0.78168 | 1.3279
| 8.2167 | 0.67358 | 3.9838 | 0.78212 | 1.3279
| 8.15 | 0.67358 | 3.984 | 0.78397 | 1.328
| 8.2 | 0.67358 | 3.9842 | 0.78257 | 1.3281
| 8.1333 | 0.67715 | 3.9844 | 0.78442 | 1.3281
| 8.1833 | 0.67358 | 3.9844 | 0.78301 | 1.3281
| 8.1667 | 0.67715 | 3.9847 | 0.78345 | 1.3282
| 8.15 | 0.67358 | 3.9851 | 0.78389 | 1.3284
| 8.1333 | 0.67358 | 3.9853 | 0.78433 | 1.3284
| 8.0667 | 0.67358 | 3.9856 | 0.78618 | 1.3285
| 8.1167 | 0.67358 | 3.9857 | 0.78478 | 1.3286
| 8.05 | 0.67715 | 3.986 | 0.78662 | 1.3287
| 8.1 | 0.67715 | 3.986 | 0.78522 | 1.3287
| 8.0333 | 0.67715 | 3.9862 | 0.78707 | 1.3287
| 8.0833 | 0.67715 | 3.9864 | 0.78566 | 1.3288
| 8.0167 | 0.67358 | 3.9866 | 0.78751 | 1.3289
| 8.0667 | 0.67358 | 3.9867 | 0.7861 | 1.3289
| 8 | 0.67715 | 3.987 | 0.78795 | 1.329
| 8.05 | 0.67358 | 3.987 | 0.78654 | 1.329
| 8.0333 | 0.67715 | 3.9872 | 0.78699 | 1.3291
| 7.9667 | 0.67358 | 3.9874 | 0.78883 | 1.3291
| 8.0167 | 0.67358 | 3.9876 | 0.78743 | 1.3292
| 7.95 | 0.67358 | 3.9877 | 0.78928 | 1.3292
| 8 | 0.67715 | 3.988 | 0.78787 | 1.3293
| 7.9333 | 0.67715 | 3.988 | 0.78972 | 1.3293
| 7.9833 | 0.67358 | 3.9882 | 0.78831 | 1.3294
| 7.9667 | 0.67358 | 3.9884 | 0.78875 | 1.3295
| 7.9167 | 0.67358 | 3.9885 | 0.79016 | 1.3295
| 7.95 | 0.67715 | 3.9889 | 0.7892 | 1.3296
| 7.8833 | 0.67358 | 3.989 | 0.79104 | 1.3297
| 7.9333 | 0.67715 | 3.9892 | 0.78964 | 1.3297
| 7.8667 | 0.67358 | 3.9893 | 0.79148 | 1.3298
| 7.9167 | 0.67715 | 3.9895 | 0.79008 | 1.3298
| 7.9 | 0.67715 | 3.9898 | 0.79052 | 1.3299
| 7.8833 | 0.67358 | 3.99 | 0.79096 | 1.33
| 7.8167 | 0.67715 | 3.9902 | 0.79281 | 1.3301
| 7.8667 | 0.67358 | 3.9904 | 0.79141 | 1.3301
| 7.8 | 0.67715 | 3.9905 | 0.79325 | 1.3302
| 7.85 | 0.67715 | 3.9906 | 0.79185 | 1.3302
| 7.7833 | 0.67358 | 3.9908 | 0.79369 | 1.3303
| 7.8333 | 0.67715 | 3.9909 | 0.79229 | 1.3303
| 7.8167 | 0.67358 | 3.9912 | 0.79273 | 1.3304
| 7.7667 | 0.67358 | 3.9912 | 0.79413 | 1.3304
| 7.75 | 0.67715 | 3.9914 | 0.79458 | 1.3305
| 7.8 | 0.67715 | 3.9916 | 0.79317 | 1.3305
| 7.7333 | 0.67715 | 3.9918 | 0.79502 | 1.3306
| 7.7833 | 0.67358 | 3.992 | 0.79361 | 1.3307
| 7.7167 | 0.67358 | 3.9921 | 0.79546 | 1.3307
| 7.7667 | 0.67358 | 3.9923 | 0.79406 | 1.3308
| 7.75 | 0.67358 | 3.9926 | 0.7945 | 1.3309
| 7.7333 | 0.67358 | 3.9927 | 0.79494 | 1.3309
| 7.6667 | 0.67358 | 3.993 | 0.79679 | 1.331
| 7.7167 | 0.67358 | 3.9931 | 0.79538 | 1.331
| 7.65 | 0.67715 | 3.9932 | 0.79723 | 1.3311
| 7.7 | 0.67358 | 3.9933 | 0.79582 | 1.3311
| 7.6833 | 0.67715 | 3.9936 | 0.79627 | 1.3312
| 7.6167 | 0.67358 | 3.9939 | 0.79811 | 1.3313
| 7.6667 | 0.67715 | 3.9941 | 0.79671 | 1.3314
| 7.6 | 0.67358 | 3.9943 | 0.79855 | 1.3314
| 7.65 | 0.67358 | 3.9945 | 0.79715 | 1.3315
| 7.5833 | 0.67001 | 3.9946 | 0.79899 | 1.3315
| 7.6333 | 0.67358 | 3.9948 | 0.79759 | 1.3316
| 7.5667 | 0.67358 | 3.9949 | 0.79944 | 1.3316
| 7.55 | 0.67358 | 3.9951 | 0.79988 | 1.3317
| 7.6167 | 0.67358 | 3.9952 | 0.79803 | 1.3317
| 7.5333 | 0.67715 | 3.9954 | 0.80032 | 1.3318
| 7.6 | 0.67358 | 3.9954 | 0.79848 | 1.3318
| 7.5833 | 0.67358 | 3.9957 | 0.79892 | 1.3319
| 7.5167 | 0.67715 | 3.9958 | 0.80076 | 1.3319
| 7.5667 | 0.67358 | 3.996 | 0.79936 | 1.332
| 7.55 | 0.67358 | 3.9962 | 0.7998 | 1.3321
| 7.4833 | 0.67358 | 3.9963 | 0.80164 | 1.3321
| 7.5333 | 0.67358 | 3.9965 | 0.80024 | 1.3322
| 7.4667 | 0.67715 | 3.9968 | 0.80209 | 1.3323
| 7.5167 | 0.67358 | 3.997 | 0.80069 | 1.3323
| 7.45 | 0.67358 | 3.997 | 0.80253 | 1.3323
| 7.5 | 0.67358 | 3.9971 | 0.80113 | 1.3324
| 7.4333 | 0.67715 | 3.9974 | 0.80297 | 1.3325
| 7.4833 | 0.67358 | 3.9975 | 0.80157 | 1.3325
| 7.4167 | 0.67358 | 3.9976 | 0.80341 | 1.3325
| 7.4667 | 0.67001 | 3.9978 | 0.80201 | 1.3326
| 7.4 | 0.67358 | 3.9979 | 0.80385 | 1.3326
| 7.3833 | 0.67358 | 3.9982 | 0.8043 | 1.3327
| 7.45 | 0.67358 | 3.9983 | 0.80245 | 1.3328
| 7.4333 | 0.67358 | 3.9985 | 0.8029 | 1.3328
| 7.3667 | 0.67715 | 3.9985 | 0.80474 | 1.3328
| 7.4167 | 0.67358 | 3.9988 | 0.80334 | 1.3329
| 7.4 | 0.67715 | 3.999 | 0.80378 | 1.333
| 7.3333 | 0.67358 | 3.9992 | 0.80562 | 1.3331
| 7.3833 | 0.67715 | 3.9994 | 0.80422 | 1.3331
| 7.3167 | 0.67358 | 3.9995 | 0.80606 | 1.3332
| 7.3667 | 0.67715 | 3.9997 | 0.80466 | 1.3332
| 7.3 | 0.67715 | 3.9998 | 0.8065 | 1.3333
| 7.35 | 0.67715 | 4 | 0.80511 | 1.3333
| 7.2833 | 0.67358 | 4.0001 | 0.80695 | 1.3334
| 7.2667 | 0.67358 | 4.0004 | 0.80739 | 1.3335
| 7.3333 | 0.67358 | 4.0005 | 0.80555 | 1.3335
| 7.3167 | 0.67358 | 4.0008 | 0.80599 | 1.3336
| 7.3 | 0.67715 | 4.0011 | 0.80643 | 1.3337
| 7.2333 | 0.67358 | 4.0011 | 0.80827 | 1.3337
| 7.2167 | 0.67358 | 4.0013 | 0.80871 | 1.3338
| 7.2833 | 0.67715 | 4.0014 | 0.80687 | 1.3338
| 7.2667 | 0.67358 | 4.0016 | 0.80731 | 1.3339
| 7.2 | 0.67715 | 4.0018 | 0.80916 | 1.3339
| 7.25 | 0.67358 | 4.0019 | 0.80776 | 1.334
| 7.1833 | 0.67001 | 4.0021 | 0.8096 | 1.334
| 7.2333 | 0.67001 | 4.0022 | 0.8082 | 1.3341
| 7.1667 | 0.67715 | 4.0024 | 0.81004 | 1.3341
| 7.2167 | 0.67358 | 4.0025 | 0.80864 | 1.3342
| 7.15 | 0.67358 | 4.0027 | 0.81048 | 1.3342
| 7.2 | 0.67358 | 4.0029 | 0.80908 | 1.3343
| 7.1333 | 0.67715 | 4.003 | 0.81092 | 1.3343
| 7.1167 | 0.67715 | 4.0033 | 0.81136 | 1.3344
| 7.1833 | 0.67358 | 4.0034 | 0.80952 | 1.3345
| 7.1 | 0.67715 | 4.0036 | 0.81181 | 1.3345
| 7.1667 | 0.67715 | 4.0037 | 0.80997 | 1.3346
| 7.0833 | 0.67358 | 4.0038 | 0.81225 | 1.3346
| 7.15 | 0.67358 | 4.0039 | 0.81041 | 1.3346
| 7.1333 | 0.67358 | 4.0041 | 0.81085 | 1.3347
| 7.0667 | 0.67358 | 4.0043 | 0.81269 | 1.3348
| 7.1167 | 0.67358 | 4.0044 | 0.81129 | 1.3348
| 7.05 | 0.67358 | 4.0045 | 0.81313 | 1.3348
| 7.1 | 0.67715 | 4.0047 | 0.81173 | 1.3349
| 7.0333 | 0.67001 | 4.0049 | 0.81357 | 1.335
| 7.0833 | 0.67001 | 4.0051 | 0.81218 | 1.335
| 7.0667 | 0.67358 | 4.0053 | 0.81262 | 1.3351
| 7 | 0.67358 | 4.0056 | 0.81446 | 1.3352
| 7.05 | 0.67715 | 4.0058 | 0.81306 | 1.3353
| 6.9833 | 0.67358 | 4.0059 | 0.8149 | 1.3353
| 7.0333 | 0.67358 | 4.0061 | 0.8135 | 1.3354
| 6.9667 | 0.67715 | 4.0063 | 0.81534 | 1.3354
| 6.95 | 0.67358 | 4.0064 | 0.81578 | 1.3355
| 7.0167 | 0.67715 | 4.0066 | 0.81394 | 1.3355
| 6.9333 | 0.67715 | 4.0067 | 0.81622 | 1.3356
| 7 | 0.67358 | 4.0068 | 0.81439 | 1.3356
| 6.9833 | 0.67715 | 4.0071 | 0.81483 | 1.3357
| 6.9667 | 0.67358 | 4.0072 | 0.81527 | 1.3357
| 6.9 | 0.67358 | 4.0074 | 0.81711 | 1.3358
| 6.95 | 0.67358 | 4.0076 | 0.81571 | 1.3359
| 6.8833 | 0.67358 | 4.0078 | 0.81755 | 1.3359
| 6.9333 | 0.67358 | 4.0081 | 0.81615 | 1.336
| 6.9167 | 0.67358 | 4.0084 | 0.8166 | 1.3361
| 6.85 | 0.67358 | 4.0085 | 0.81843 | 1.3362
| 6.9 | 0.67358 | 4.0087 | 0.81704 | 1.3362
| 6.8833 | 0.67358 | 4.009 | 0.81748 | 1.3363
| 6.8667 | 0.67715 | 4.0093 | 0.81792 | 1.3364
| 6.8 | 0.67715 | 4.0094 | 0.81976 | 1.3365
| 6.85 | 0.67715 | 4.0096 | 0.81836 | 1.3365
| 6.8333 | 0.67358 | 4.0098 | 0.81881 | 1.3366
| 6.7667 | 0.67715 | 4.01 | 0.82064 | 1.3367
| 6.8167 | 0.67358 | 4.0101 | 0.81925 | 1.3367
| 6.75 | 0.67715 | 4.0102 | 0.82108 | 1.3367
| 6.8 | 0.67715 | 4.0105 | 0.81969 | 1.3368
| 6.7333 | 0.67358 | 4.0106 | 0.82152 | 1.3369
| 6.7833 | 0.67715 | 4.0108 | 0.82013 | 1.3369
| 6.7167 | 0.67715 | 4.0111 | 0.82197 | 1.337
| 6.7 | 0.67358 | 4.0112 | 0.82241 | 1.3371
| 6.7667 | 0.67715 | 4.0113 | 0.82057 | 1.3371
| 6.75 | 0.67715 | 4.0117 | 0.82101 | 1.3372
| 6.6833 | 0.67715 | 4.0118 | 0.82285 | 1.3373
| 6.7333 | 0.67001 | 4.0119 | 0.82146 | 1.3373
| 6.65 | 0.67001 | 4.0122 | 0.82373 | 1.3374
| 6.7167 | 0.67715 | 4.0123 | 0.8219 | 1.3374
| 6.6333 | 0.67715 | 4.0124 | 0.82418 | 1.3375
| 6.7 | 0.67715 | 4.0126 | 0.82234 | 1.3375
| 6.6833 | 0.67358 | 4.0129 | 0.82278 | 1.3376
| 6.6 | 0.67358 | 4.0131 | 0.82506 | 1.3377
| 6.6667 | 0.67001 | 4.0132 | 0.82322 | 1.3377
| 6.65 | 0.67715 | 4.0134 | 0.82367 | 1.3378
| 6.5833 | 0.67715 | 4.0137 | 0.8255 | 1.3379
| 6.6333 | 0.67358 | 4.0138 | 0.82411 | 1.3379
| 6.5667 | 0.67715 | 4.014 | 0.82594 | 1.338
| 6.6167 | 0.67358 | 4.0141 | 0.82455 | 1.338
| 6.55 | 0.68072 | 4.0143 | 0.82638 | 1.3381
| 6.6 | 0.67358 | 4.0145 | 0.82499 | 1.3382
| 6.5833 | 0.67358 | 4.0148 | 0.82543 | 1.3383
| 6.5167 | 0.67715 | 4.0149 | 0.82727 | 1.3383
| 6.5 | 0.67715 | 4.015 | 0.82771 | 1.3383
| 6.5667 | 0.67358 | 4.0153 | 0.82588 | 1.3384
| 6.4833 | 0.67358 | 4.0154 | 0.82815 | 1.3385
| 6.55 | 0.67715 | 4.0156 | 0.82632 | 1.3385
| 6.4667 | 0.67358 | 4.0159 | 0.82859 | 1.3386
| 6.5333 | 0.67358 | 4.016 | 0.82676 | 1.3387
| 6.45 | 0.67358 | 4.0161 | 0.82903 | 1.3387
| 6.5167 | 0.67715 | 4.0162 | 0.8272 | 1.3387
| 6.5 | 0.67715 | 4.0164 | 0.82764 | 1.3388
| 6.4833 | 0.67358 | 4.0167 | 0.82809 | 1.3389
| 6.4167 | 0.67358 | 4.0168 | 0.82992 | 1.3389
| 6.4667 | 0.67358 | 4.0171 | 0.82853 | 1.339
| 6.4 | 0.67358 | 4.0172 | 0.83036 | 1.3391
| 6.3833 | 0.67358 | 4.0174 | 0.8308 | 1.3391
| 6.45 | 0.67358 | 4.0176 | 0.82897 | 1.3392
| 6.3667 | 0.67715 | 4.0179 | 0.83124 | 1.3393
| 6.4333 | 0.67715 | 4.0179 | 0.82941 | 1.3393
| 6.35 | 0.67358 | 4.0181 | 0.83169 | 1.3394
| 6.4167 | 0.67715 | 4.0184 | 0.82985 | 1.3395
| 6.4 | 0.67358 | 4.0185 | 0.8303 | 1.3395
| 6.3833 | 0.67358 | 4.0186 | 0.83074 | 1.3395
| 6.3167 | 0.67358 | 4.0188 | 0.83257 | 1.3396
| 6.3667 | 0.67715 | 4.019 | 0.83118 | 1.3397
| 6.3 | 0.67715 | 4.0191 | 0.83301 | 1.3397
| 6.2833 | 0.67358 | 4.0194 | 0.83345 | 1.3398
| 6.35 | 0.67358 | 4.0195 | 0.83162 | 1.3398
| 6.3333 | 0.67358 | 4.0198 | 0.83206 | 1.3399
| 6.2667 | 0.67715 | 4.0199 | 0.83389 | 1.34
| 6.3167 | 0.67358 | 4.0201 | 0.83251 | 1.34
| 6.3 | 0.67715 | 4.0205 | 0.83295 | 1.3402
| 6.2333 | 0.67358 | 4.0206 | 0.83478 | 1.3402
| 6.2167 | 0.67358 | 4.0208 | 0.83522 | 1.3403
| 6.2833 | 0.67358 | 4.021 | 0.83339 | 1.3403
| 6.2667 | 0.67358 | 4.0212 | 0.83383 | 1.3404
| 6.1833 | 0.67358 | 4.0214 | 0.8361 | 1.3405
| 6.25 | 0.67715 | 4.0215 | 0.83427 | 1.3405
| 6.1667 | 0.67358 | 4.0217 | 0.83654 | 1.3406
| 6.2333 | 0.67358 | 4.0219 | 0.83472 | 1.3406
| 6.2167 | 0.67715 | 4.0221 | 0.83516 | 1.3407
| 6.1333 | 0.67358 | 4.0224 | 0.83743 | 1.3408
| 6.2 | 0.67358 | 4.0226 | 0.8356 | 1.3409
| 6.1833 | 0.67715 | 4.0227 | 0.83604 | 1.3409
| 6.1 | 0.67715 | 4.0231 | 0.83831 | 1.341
| 6.1667 | 0.67358 | 4.0232 | 0.83648 | 1.3411
| 6.15 | 0.67715 | 4.0234 | 0.83692 | 1.3411
| 6.0833 | 0.67715 | 4.0235 | 0.83875 | 1.3412
| 6.1333 | 0.67715 | 4.0238 | 0.83737 | 1.3413
| 6.05 | 0.67001 | 4.0242 | 0.83964 | 1.3414
| 6.1167 | 0.67715 | 4.0243 | 0.83781 | 1.3414
| 6.0333 | 0.67715 | 4.0244 | 0.84008 | 1.3415
| 6.1 | 0.67715 | 4.0246 | 0.83825 | 1.3415
| 6.0167 | 0.67358 | 4.0248 | 0.84052 | 1.3416
| 6.0833 | 0.67358 | 4.0249 | 0.83869 | 1.3416
| 6.0667 | 0.67715 | 4.025 | 0.83913 | 1.3417
| 5.9833 | 0.67715 | 4.0253 | 0.8414 | 1.3418
| 6.05 | 0.67358 | 4.0255 | 0.83958 | 1.3418
| 6.0333 | 0.67358 | 4.0259 | 0.84002 | 1.342
| 6.0167 | 0.67358 | 4.026 | 0.84046 | 1.342
| 5.95 | 0.67358 | 4.0262 | 0.84229 | 1.3421
| 5.9333 | 0.67715 | 4.0266 | 0.84273 | 1.3422
| 6 | 0.67001 | 4.0266 | 0.8409 | 1.3422
| 5.9167 | 0.67358 | 4.0268 | 0.84317 | 1.3423
| 5.9833 | 0.67715 | 4.0269 | 0.84134 | 1.3423
| 5.9 | 0.67358 | 4.0272 | 0.84361 | 1.3424
| 5.9667 | 0.67715 | 4.0273 | 0.84179 | 1.3424
| 5.8833 | 0.67358 | 4.0275 | 0.84405 | 1.3425
| 5.95 | 0.67715 | 4.0275 | 0.84223 | 1.3425
| 5.8667 | 0.67001 | 4.0277 | 0.8445 | 1.3426
| 5.9333 | 0.67001 | 4.0278 | 0.84267 | 1.3426
| 5.85 | 0.67358 | 4.0282 | 0.84494 | 1.3427
| 5.9167 | 0.67358 | 4.0283 | 0.84311 | 1.3428
| 5.9 | 0.67358 | 4.0285 | 0.84355 | 1.3428
| 5.8333 | 0.67715 | 4.0285 | 0.84538 | 1.3428
| 5.8167 | 0.67715 | 4.0288 | 0.84582 | 1.3429
| 5.8833 | 0.67715 | 4.0289 | 0.844 | 1.343
| 5.8667 | 0.67358 | 4.0292 | 0.84444 | 1.3431
| 5.85 | 0.67358 | 4.0295 | 0.84488 | 1.3432
| 5.7667 | 0.67358 | 4.0298 | 0.84715 | 1.3433
| 5.8333 | 0.67715 | 4.03 | 0.84532 | 1.3433
| 5.75 | 0.67358 | 4.0301 | 0.84759 | 1.3434
| 5.8167 | 0.67715 | 4.0302 | 0.84576 | 1.3434
| 5.7333 | 0.67358 | 4.0304 | 0.84803 | 1.3435
| 5.8 | 0.67358 | 4.0306 | 0.84621 | 1.3435
| 5.7167 | 0.67715 | 4.0308 | 0.84847 | 1.3436
| 5.7833 | 0.67358 | 4.0309 | 0.84665 | 1.3436
| 5.7 | 0.67358 | 4.031 | 0.84891 | 1.3437
| 5.7667 | 0.67715 | 4.0313 | 0.84709 | 1.3438
| 5.6833 | 0.67715 | 4.0315 | 0.84936 | 1.3438
| 5.75 | 0.67358 | 4.0317 | 0.84753 | 1.3439
| 5.6667 | 0.67358 | 4.0318 | 0.8498 | 1.3439
| 5.7333 | 0.67358 | 4.032 | 0.84797 | 1.344
| 5.7167 | 0.67358 | 4.0322 | 0.84842 | 1.3441
| 5.6333 | 0.67358 | 4.0326 | 0.85068 | 1.3442
| 5.7 | 0.67715 | 4.0327 | 0.84886 | 1.3442
| 5.6167 | 0.67358 | 4.0328 | 0.85112 | 1.3443
| 5.6833 | 0.67358 | 4.0331 | 0.8493 | 1.3444
| 5.6667 | 0.67715 | 4.0333 | 0.84974 | 1.3444
| 5.5833 | 0.67715 | 4.0336 | 0.85201 | 1.3445
| 5.65 | 0.67358 | 4.0337 | 0.85018 | 1.3446
| 5.5667 | 0.67358 | 4.0338 | 0.85245 | 1.3446
| 5.6333 | 0.67715 | 4.0341 | 0.85062 | 1.3447
| 5.55 | 0.67715 | 4.0342 | 0.85289 | 1.3447
| 5.6167 | 0.67715 | 4.0344 | 0.85107 | 1.3448
| 5.5333 | 0.67358 | 4.0346 | 0.85333 | 1.3449
| 5.6 | 0.67358 | 4.0346 | 0.85151 | 1.3449
| 5.5167 | 0.67358 | 4.0349 | 0.85377 | 1.345
| 5.5833 | 0.67358 | 4.0349 | 0.85195 | 1.345
| 5.5667 | 0.67715 | 4.0352 | 0.85239 | 1.3451
| 5.5 | 0.67358 | 4.0353 | 0.85422 | 1.3451
| 5.4833 | 0.67358 | 4.0355 | 0.85466 | 1.3452
| 5.55 | 0.67715 | 4.0357 | 0.85283 | 1.3452
| 5.4667 | 0.67715 | 4.0359 | 0.8551 | 1.3453
| 5.5333 | 0.67715 | 4.036 | 0.85328 | 1.3453
| 5.45 | 0.67358 | 4.0362 | 0.85554 | 1.3454
| 5.5167 | 0.67358 | 4.0365 | 0.85372 | 1.3455
| 5.4333 | 0.67715 | 4.0366 | 0.85598 | 1.3455
| 5.5 | 0.67715 | 4.0367 | 0.85416 | 1.3456
| 5.4833 | 0.67715 | 4.0371 | 0.8546 | 1.3457
| 5.4 | 0.67358 | 4.0373 | 0.85687 | 1.3458
| 5.4667 | 0.67715 | 4.0374 | 0.85504 | 1.3458
| 5.45 | 0.67358 | 4.0378 | 0.85549 | 1.3459
| 5.3667 | 0.67358 | 4.0379 | 0.85775 | 1.346
| 5.4333 | 0.67358 | 4.0381 | 0.85593 | 1.346
| 5.35 | 0.67715 | 4.0383 | 0.85819 | 1.3461
| 5.4167 | 0.67358 | 4.0385 | 0.85637 | 1.3461
| 5.3333 | 0.67358 | 4.0386 | 0.85863 | 1.3462
| 5.4 | 0.67001 | 4.0388 | 0.85681 | 1.3463
| 5.3167 | 0.67358 | 4.039 | 0.85908 | 1.3463
| 5.3833 | 0.67358 | 4.0391 | 0.85725 | 1.3464
| 5.3 | 0.67358 | 4.0393 | 0.85952 | 1.3464
| 5.3667 | 0.67358 | 4.0395 | 0.8577 | 1.3465
| 5.2833 | 0.67358 | 4.0396 | 0.85996 | 1.3465
| 5.35 | 0.67358 | 4.0398 | 0.85814 | 1.3466
| 5.2667 | 0.67358 | 4.04 | 0.8604 | 1.3467
| 5.3333 | 0.67358 | 4.0402 | 0.85858 | 1.3467
| 5.25 | 0.67358 | 4.0403 | 0.86084 | 1.3468
| 5.3167 | 0.67715 | 4.0405 | 0.85902 | 1.3468
| 5.2333 | 0.67001 | 4.0407 | 0.86128 | 1.3469
| 5.3 | 0.67715 | 4.041 | 0.85946 | 1.347
| 5.2167 | 0.67358 | 4.0411 | 0.86173 | 1.347
| 5.2833 | 0.67715 | 4.0413 | 0.85991 | 1.3471
| 5.2667 | 0.67715 | 4.0414 | 0.86035 | 1.3471
| 5.2 | 0.67715 | 4.0415 | 0.86217 | 1.3472
| 5.1833 | 0.67001 | 4.0417 | 0.86261 | 1.3472
| 5.25 | 0.67358 | 4.0418 | 0.86079 | 1.3473
| 5.1667 | 0.67715 | 4.0422 | 0.86305 | 1.3474
| 5.2333 | 0.67358 | 4.0423 | 0.86123 | 1.3474
| 5.15 | 0.67001 | 4.0424 | 0.86349 | 1.3475
| 5.2167 | 0.67715 | 4.0425 | 0.86167 | 1.3475
| 5.1333 | 0.67358 | 4.0429 | 0.86393 | 1.3476
| 5.2 | 0.67358 | 4.043 | 0.86212 | 1.3477
| 5.1167 | 0.67358 | 4.0431 | 0.86438 | 1.3477
| 5.1833 | 0.67715 | 4.0434 | 0.86256 | 1.3478
| 5.1 | 0.67715 | 4.0435 | 0.86482 | 1.3478
| 5.1667 | 0.67715 | 4.0436 | 0.863 | 1.3479
| 5.0833 | 0.68072 | 4.0439 | 0.86526 | 1.348
| 5.15 | 0.67715 | 4.044 | 0.86344 | 1.348
| 5.0667 | 0.67715 | 4.0442 | 0.8657 | 1.3481
| 5.1333 | 0.67358 | 4.0443 | 0.86388 | 1.3481
| 5.1167 | 0.67358 | 4.0447 | 0.86432 | 1.3482
| 5.1 | 0.67715 | 4.045 | 0.86477 | 1.3483
| 5.0167 | 0.67715 | 4.0453 | 0.86703 | 1.3484
| 5.0833 | 0.67715 | 4.0455 | 0.86521 | 1.3485
| 5 | 0.67715 | 4.0456 | 0.86747 | 1.3485
| 5.0667 | 0.67358 | 4.0459 | 0.86565 | 1.3486
| 5.05 | 0.67358 | 4.0462 | 0.86609 | 1.3487
| 4.95 | 0.67358 | 4.0465 | 0.86879 | 1.3488
| 5.0333 | 0.67358 | 4.0465 | 0.86653 | 1.3488
| 5.0167 | 0.67715 | 4.0468 | 0.86698 | 1.3489
| 4.9333 | 0.67358 | 4.047 | 0.86924 | 1.349
| 5 | 0.67715 | 4.0472 | 0.86742 | 1.3491
| 4.9833 | 0.67358 | 4.0474 | 0.86786 | 1.3491
| 4.9 | 0.67715 | 4.0477 | 0.87012 | 1.3492
| 4.9667 | 0.67358 | 4.0478 | 0.8683 | 1.3493
| 4.8833 | 0.67358 | 4.0481 | 0.87056 | 1.3494
| 4.95 | 0.67715 | 4.0482 | 0.86874 | 1.3494
| 4.8667 | 0.67358 | 4.0483 | 0.871 | 1.3494
| 4.85 | 0.68072 | 4.0486 | 0.87144 | 1.3495
| 4.9333 | 0.67001 | 4.0487 | 0.86919 | 1.3496
| 4.9167 | 0.67358 | 4.0489 | 0.86963 | 1.3496
| 4.8333 | 0.68072 | 4.0491 | 0.87189 | 1.3497
| 4.9 | 0.68072 | 4.0493 | 0.87007 | 1.3498
| 4.8167 | 0.67358 | 4.0494 | 0.87233 | 1.3498
| 4.8833 | 0.67358 | 4.0496 | 0.87051 | 1.3499
| 4.8 | 0.67715 | 4.0497 | 0.87277 | 1.3499
| 4.8667 | 0.67358 | 4.0499 | 0.87095 | 1.35
| 4.85 | 0.67358 | 4.0502 | 0.8714 | 1.3501
| 4.7667 | 0.67358 | 4.0504 | 0.87365 | 1.3501
| 4.8333 | 0.67715 | 4.0506 | 0.87184 | 1.3502
| 4.75 | 0.67358 | 4.0508 | 0.8741 | 1.3503
| 4.8167 | 0.67358 | 4.0509 | 0.87228 | 1.3503
| 4.7333 | 0.67358 | 4.0512 | 0.87454 | 1.3504
| 4.8 | 0.67715 | 4.0513 | 0.87272 | 1.3504
| 4.7167 | 0.67358 | 4.0514 | 0.87498 | 1.3505
| 4.7833 | 0.67358 | 4.0517 | 0.87316 | 1.3506
| 4.7 | 0.67358 | 4.0518 | 0.87542 | 1.3506
| 4.7667 | 0.67358 | 4.0521 | 0.87361 | 1.3507
| 4.6833 | 0.67715 | 4.0522 | 0.87586 | 1.3507
| 4.75 | 0.67358 | 4.0523 | 0.87405 | 1.3508
| 4.6667 | 0.67358 | 4.0525 | 0.8763 | 1.3508
| 4.7333 | 0.67715 | 4.0526 | 0.87449 | 1.3509
| 4.65 | 0.67358 | 4.0529 | 0.87675 | 1.351
| 4.7167 | 0.67715 | 4.053 | 0.87493 | 1.351
| 4.6333 | 0.67358 | 4.0533 | 0.87719 | 1.3511
| 4.7 | 0.67358 | 4.0533 | 0.87537 | 1.3511
| 4.6167 | 0.67715 | 4.0535 | 0.87763 | 1.3512
| 4.6833 | 0.67358 | 4.0536 | 0.87582 | 1.3512
| 4.6 | 0.67358 | 4.0539 | 0.87807 | 1.3513
| 4.6667 | 0.67358 | 4.054 | 0.87626 | 1.3513
| 4.5833 | 0.67358 | 4.0542 | 0.87851 | 1.3514
| 4.65 | 0.67358 | 4.0544 | 0.8767 | 1.3515
| 4.5667 | 0.67358 | 4.0546 | 0.87895 | 1.3515
| 4.6333 | 0.67715 | 4.0549 | 0.87714 | 1.3516
| 4.6167 | 0.67358 | 4.0552 | 0.87758 | 1.3517
| 4.6 | 0.67715 | 4.0555 | 0.87802 | 1.3518
| 4.5167 | 0.67358 | 4.0556 | 0.88028 | 1.3519
| 4.5833 | 0.67358 | 4.0558 | 0.87847 | 1.3519
| 4.5 | 0.67358 | 4.056 | 0.88072 | 1.352
| 4.5667 | 0.67358 | 4.0563 | 0.87891 | 1.3521
| 4.55 | 0.67715 | 4.0565 | 0.87935 | 1.3522
| 4.4667 | 0.67001 | 4.0567 | 0.88161 | 1.3522
| 4.5333 | 0.67358 | 4.0568 | 0.87979 | 1.3523
| 4.45 | 0.67358 | 4.0571 | 0.88205 | 1.3524
| 4.5167 | 0.67358 | 4.0572 | 0.88023 | 1.3524
| 4.4333 | 0.67715 | 4.0575 | 0.88249 | 1.3525
| 4.5 | 0.67358 | 4.0575 | 0.88068 | 1.3525
| 4.4167 | 0.67715 | 4.0578 | 0.88293 | 1.3526
| 4.4833 | 0.67358 | 4.058 | 0.88112 | 1.3527
| 4.4 | 0.67001 | 4.0581 | 0.88337 | 1.3527
| 4.3833 | 0.67715 | 4.0583 | 0.88381 | 1.3528
| 4.4667 | 0.67358 | 4.0584 | 0.88156 | 1.3528
| 4.3667 | 0.67358 | 4.0587 | 0.88426 | 1.3529
| 4.45 | 0.67715 | 4.0587 | 0.882 | 1.3529
| 4.4333 | 0.67001 | 4.0591 | 0.88244 | 1.353
| 4.35 | 0.67358 | 4.0591 | 0.8847 | 1.353
| 4.4167 | 0.67001 | 4.0594 | 0.88289 | 1.3531
| 4.3333 | 0.67715 | 4.0595 | 0.88514 | 1.3532
| 4.4 | 0.68072 | 4.0597 | 0.88333 | 1.3532
| 4.3167 | 0.67715 | 4.06 | 0.88558 | 1.3533
| 4.3833 | 0.67715 | 4.06 | 0.88377 | 1.3533
| 4.3 | 0.67358 | 4.0601 | 0.88602 | 1.3534
| 4.3667 | 0.67001 | 4.0603 | 0.88421 | 1.3534
| 4.2833 | 0.67358 | 4.0605 | 0.88646 | 1.3535
| 4.35 | 0.67358 | 4.0606 | 0.88465 | 1.3536
| 4.2667 | 0.67358 | 4.0607 | 0.88691 | 1.3536
| 4.3333 | 0.67358 | 4.0611 | 0.8851 | 1.3537
| 4.25 | 0.67715 | 4.0612 | 0.88735 | 1.3537
| 4.3167 | 0.67358 | 4.0614 | 0.88554 | 1.3538
| 4.2333 | 0.68072 | 4.0616 | 0.88779 | 1.3539
| 4.2167 | 0.67358 | 4.0618 | 0.88823 | 1.3539
| 4.3 | 0.67358 | 4.0619 | 0.88598 | 1.354
| 4.2833 | 0.67715 | 4.0621 | 0.88642 | 1.354
| 4.2 | 0.67358 | 4.0623 | 0.88867 | 1.3541
| 4.2667 | 0.67715 | 4.0625 | 0.88686 | 1.3542
| 4.1833 | 0.67715 | 4.0626 | 0.88912 | 1.3542
| 4.25 | 0.67358 | 4.0629 | 0.88731 | 1.3543
| 4.1667 | 0.67358 | 4.063 | 0.88956 | 1.3543
| 4.2333 | 0.67715 | 4.0633 | 0.88775 | 1.3544
| 4.15 | 0.67358 | 4.0633 | 0.89 | 1.3544
| 4.2167 | 0.67358 | 4.0636 | 0.88819 | 1.3545
| 4.2 | 0.67001 | 4.0638 | 0.88863 | 1.3546
| 4.1167 | 0.67715 | 4.064 | 0.89088 | 1.3547
| 4.1833 | 0.67358 | 4.0643 | 0.88907 | 1.3548
| 4.1 | 0.67358 | 4.0644 | 0.89132 | 1.3548
| 4.1667 | 0.67358 | 4.0646 | 0.88952 | 1.3549
| 4.15 | 0.67358 | 4.065 | 0.88996 | 1.355
| 4.0667 | 0.67715 | 4.0651 | 0.89221 | 1.355
| 4.1333 | 0.67715 | 4.0653 | 0.8904 | 1.3551
| 4.05 | 0.67358 | 4.0655 | 0.89265 | 1.3552
| 4.1167 | 0.67358 | 4.0656 | 0.89084 | 1.3552
| 4.1 | 0.67715 | 4.066 | 0.89128 | 1.3553
| 4.0833 | 0.67715 | 4.0663 | 0.89172 | 1.3554
| 4 | 0.67715 | 4.0665 | 0.89397 | 1.3555
| 4.0667 | 0.67715 | 4.0667 | 0.89217 | 1.3556
| 3.9833 | 0.67358 | 4.0668 | 0.89442 | 1.3556
| 4.05 | 0.67715 | 4.0671 | 0.89261 | 1.3557
| 3.9667 | 0.67358 | 4.0671 | 0.89486 | 1.3557
| 3.95 | 0.67358 | 4.0674 | 0.8953 | 1.3558
| 4.0333 | 0.67358 | 4.0675 | 0.89305 | 1.3558
| 4.0167 | 0.68072 | 4.0678 | 0.89349 | 1.3559
| 3.9333 | 0.67358 | 4.0679 | 0.89574 | 1.356
| 4 | 0.67358 | 4.0682 | 0.89393 | 1.3561
| 3.9833 | 0.67358 | 4.0685 | 0.89438 | 1.3562
| 3.9 | 0.67358 | 4.0686 | 0.89663 | 1.3562
| 3.9667 | 0.67358 | 4.0688 | 0.89482 | 1.3563
| 3.8833 | 0.67715 | 4.069 | 0.89707 | 1.3563
| 3.95 | 0.67358 | 4.0693 | 0.89526 | 1.3564
| 3.8667 | 0.67358 | 4.0694 | 0.89751 | 1.3564
| 3.9333 | 0.67358 | 4.0696 | 0.8957 | 1.3565
| 3.85 | 0.67358 | 4.0697 | 0.89795 | 1.3566
| 3.9167 | 0.67358 | 4.0699 | 0.89614 | 1.3566
| 3.8333 | 0.67358 | 4.0701 | 0.89839 | 1.3567
| 3.9 | 0.67358 | 4.0703 | 0.89659 | 1.3568
| 3.8167 | 0.67358 | 4.0704 | 0.89883 | 1.3568
| 3.8833 | 0.67358 | 4.0706 | 0.89703 | 1.3569
| 3.8 | 0.67358 | 4.0707 | 0.89928 | 1.3569
| 3.7833 | 0.67358 | 4.071 | 0.89972 | 1.357
| 3.8667 | 0.67358 | 4.0712 | 0.89747 | 1.3571
| 3.85 | 0.67715 | 4.0714 | 0.89791 | 1.3571
| 3.8333 | 0.67358 | 4.0718 | 0.89835 | 1.3573
| 3.8167 | 0.67358 | 4.072 | 0.8988 | 1.3573
| 3.7333 | 0.67715 | 4.0723 | 0.90104 | 1.3574
| 3.8 | 0.67358 | 4.0723 | 0.89924 | 1.3574
| 3.7167 | 0.67715 | 4.0725 | 0.90148 | 1.3575
| 3.7833 | 0.67715 | 4.0727 | 0.89968 | 1.3576
| 3.7 | 0.67358 | 4.0729 | 0.90193 | 1.3576
| 3.7667 | 0.67001 | 4.0731 | 0.90012 | 1.3577
| 3.6833 | 0.67358 | 4.0732 | 0.90237 | 1.3577
| 3.6667 | 0.67715 | 4.0736 | 0.90281 | 1.3579
| 3.75 | 0.67358 | 4.0736 | 0.90056 | 1.3579
| 3.7333 | 0.67358 | 4.0739 | 0.90101 | 1.358
| 3.65 | 0.67358 | 4.074 | 0.90325 | 1.358
| 3.6333 | 0.67715 | 4.0742 | 0.90369 | 1.3581
| 3.7167 | 0.67358 | 4.0743 | 0.90145 | 1.3581
| 3.7 | 0.67715 | 4.0746 | 0.90189 | 1.3582
| 3.6833 | 0.67358 | 4.075 | 0.90233 | 1.3583
| 3.6 | 0.67358 | 4.0751 | 0.90458 | 1.3584
| 3.6667 | 0.67715 | 4.0752 | 0.90277 | 1.3584
| 3.5833 | 0.67358 | 4.0754 | 0.90502 | 1.3585
| 3.65 | 0.67715 | 4.0755 | 0.90322 | 1.3585
| 3.5667 | 0.67358 | 4.0757 | 0.90546 | 1.3586
| 3.6333 | 0.67358 | 4.076 | 0.90366 | 1.3587
| 3.55 | 0.67358 | 4.0762 | 0.9059 | 1.3587
| 3.6167 | 0.67715 | 4.0765 | 0.9041 | 1.3588
| 3.6 | 0.67358 | 4.0769 | 0.90454 | 1.359
| 3.5833 | 0.67715 | 4.0772 | 0.90498 | 1.3591
| 3.5667 | 0.67358 | 4.0776 | 0.90543 | 1.3592
| 3.55 | 0.67715 | 4.0779 | 0.90587 | 1.3593
| 3.4667 | 0.67358 | 4.078 | 0.90811 | 1.3593
| 3.5333 | 0.67358 | 4.0782 | 0.90631 | 1.3594
| 3.45 | 0.67715 | 4.0783 | 0.90855 | 1.3594
| 3.5167 | 0.67715 | 4.0787 | 0.90675 | 1.3596
| 3.5 | 0.67358 | 4.079 | 0.90719 | 1.3597
| 3.4167 | 0.67715 | 4.0791 | 0.90944 | 1.3597
| 3.4833 | 0.67358 | 4.0793 | 0.90763 | 1.3598
| 3.4 | 0.67715 | 4.0794 | 0.90988 | 1.3598
| 3.3833 | 0.67715 | 4.0796 | 0.91032 | 1.3599
| 3.4667 | 0.67715 | 4.0797 | 0.90808 | 1.3599
| 3.45 | 0.67358 | 4.08 | 0.90852 | 1.36
| 3.3667 | 0.67715 | 4.0801 | 0.91076 | 1.36
| 3.35 | 0.67358 | 4.0804 | 0.9112 | 1.3601
| 3.4333 | 0.67358 | 4.0805 | 0.90896 | 1.3602
| 3.4167 | 0.67358 | 4.0809 | 0.9094 | 1.3603
| 3.4 | 0.67715 | 4.0812 | 0.90984 | 1.3604
| 3.3833 | 0.67358 | 4.0816 | 0.91029 | 1.3605
| 3.3667 | 0.67358 | 4.0819 | 0.91073 | 1.3606
| 3.35 | 0.67715 | 4.0822 | 0.91117 | 1.3607
| 3.2667 | 0.67715 | 4.0823 | 0.91341 | 1.3608
| 3.3333 | 0.67358 | 4.0825 | 0.91161 | 1.3608
| 3.25 | 0.67358 | 4.0826 | 0.91385 | 1.3609
| 3.2333 | 0.67715 | 4.0829 | 0.9143 | 1.361
| 3.3167 | 0.67358 | 4.0831 | 0.91205 | 1.361
| 3.2167 | 0.67358 | 4.0833 | 0.91474 | 1.3611
| 3.3 | 0.67358 | 4.0835 | 0.9125 | 1.3612
| 3.2 | 0.67715 | 4.0836 | 0.91518 | 1.3612
| 3.2833 | 0.67358 | 4.0838 | 0.91294 | 1.3613
| 3.1833 | 0.67358 | 4.0841 | 0.91562 | 1.3614
| 3.2667 | 0.67715 | 4.0842 | 0.91338 | 1.3614
| 3.25 | 0.67358 | 4.0845 | 0.91382 | 1.3615
| 3.2333 | 0.67358 | 4.0848 | 0.91426 | 1.3616
| 3.2167 | 0.67358 | 4.0852 | 0.91471 | 1.3617
| 3.2 | 0.67358 | 4.0855 | 0.91515 | 1.3618
| 3.1167 | 0.67358 | 4.0857 | 0.91739 | 1.3619
| 3.1833 | 0.67358 | 4.0858 | 0.91559 | 1.3619
| 3.1 | 0.67358 | 4.086 | 0.91783 | 1.362
| 3.1667 | 0.67715 | 4.0864 | 0.91603 | 1.3621
| 3.15 | 0.67715 | 4.0867 | 0.91647 | 1.3622
| 3.05 | 0.67715 | 4.087 | 0.91916 | 1.3623
| 3.1333 | 0.67715 | 4.0871 | 0.91692 | 1.3624
| 3.1167 | 0.67715 | 4.0874 | 0.91736 | 1.3625
| 3.0167 | 0.67715 | 4.0877 | 0.92004 | 1.3626
| 3.1 | 0.67358 | 4.0878 | 0.9178 | 1.3626
| 3 | 0.67715 | 4.088 | 0.92048 | 1.3627
| 3.0833 | 0.67358 | 4.0881 | 0.91824 | 1.3627
| 3.0667 | 0.67715 | 4.0884 | 0.91868 | 1.3628
| 2.9833 | 0.67715 | 4.0886 | 0.92092 | 1.3629
| 3.05 | 0.67358 | 4.0889 | 0.91913 | 1.363
| 2.9667 | 0.67715 | 4.089 | 0.92136 | 1.363
| 3.0333 | 0.67715 | 4.0892 | 0.91957 | 1.3631
| 2.95 | 0.67358 | 4.0895 | 0.92181 | 1.3632
| 3.0167 | 0.67715 | 4.0896 | 0.92001 | 1.3632
| 2.9333 | 0.67358 | 4.0897 | 0.92225 | 1.3632
| 3 | 0.67358 | 4.09 | 0.92045 | 1.3633
| 2.9167 | 0.67715 | 4.0901 | 0.92269 | 1.3634
| 2.9 | 0.67715 | 4.0903 | 0.92313 | 1.3634
| 2.9833 | 0.67715 | 4.0904 | 0.92089 | 1.3635
| 2.9667 | 0.68072 | 4.0907 | 0.92133 | 1.3636
| 2.8833 | 0.67358 | 4.0909 | 0.92357 | 1.3636
| 2.8667 | 0.67358 | 4.0911 | 0.92402 | 1.3637
| 2.95 | 0.67358 | 4.0912 | 0.92178 | 1.3637
| 2.9333 | 0.67715 | 4.0914 | 0.92222 | 1.3638
| 2.85 | 0.67715 | 4.0915 | 0.92446 | 1.3639
| 2.8333 | 0.67358 | 4.0918 | 0.9249 | 1.3639
| 2.9167 | 0.67358 | 4.0919 | 0.92266 | 1.364
| 2.9 | 0.67715 | 4.0921 | 0.9231 | 1.364
| 2.8167 | 0.67715 | 4.0922 | 0.92534 | 1.3641
| 2.8833 | 0.67001 | 4.0927 | 0.92354 | 1.3642
| 2.8667 | 0.67358 | 4.093 | 0.92399 | 1.3643
| 2.7833 | 0.68072 | 4.0931 | 0.92622 | 1.3644
| 2.85 | 0.67715 | 4.0933 | 0.92443 | 1.3644
| 2.7667 | 0.67358 | 4.0934 | 0.92667 | 1.3645
| 2.75 | 0.67358 | 4.0936 | 0.92711 | 1.3645
| 2.8333 | 0.67715 | 4.0938 | 0.92487 | 1.3646
| 2.8167 | 0.67358 | 4.0941 | 0.92531 | 1.3647
| 2.8 | 0.67715 | 4.0945 | 0.92575 | 1.3648
| 2.7833 | 0.67358 | 4.0948 | 0.9262 | 1.3649
| 2.7 | 0.67358 | 4.0949 | 0.92843 | 1.365
| 2.7667 | 0.67715 | 4.0953 | 0.92664 | 1.3651
| 2.6833 | 0.68072 | 4.0954 | 0.92887 | 1.3651
| 2.75 | 0.67715 | 4.0956 | 0.92708 | 1.3652
| 2.6667 | 0.67715 | 4.0957 | 0.92932 | 1.3652
| 2.7333 | 0.67715 | 4.096 | 0.92752 | 1.3653
| 2.65 | 0.67715 | 4.0961 | 0.92976 | 1.3654
| 2.7167 | 0.67358 | 4.0964 | 0.92796 | 1.3655
| 2.7 | 0.67358 | 4.0967 | 0.92841 | 1.3656
| 2.6167 | 0.67358 | 4.0968 | 0.93064 | 1.3656
| 2.6 | 0.67358 | 4.0971 | 0.93108 | 1.3657
| 2.6833 | 0.67358 | 4.0973 | 0.92885 | 1.3658
| 2.5833 | 0.67358 | 4.0974 | 0.93153 | 1.3658
| 2.6667 | 0.67358 | 4.0977 | 0.92929 | 1.3659
| 2.5667 | 0.67358 | 4.0978 | 0.93197 | 1.3659
| 2.65 | 0.67715 | 4.0979 | 0.92973 | 1.366
| 2.6333 | 0.67715 | 4.0983 | 0.93017 | 1.3661
| 2.55 | 0.67715 | 4.0983 | 0.93241 | 1.3661
| 2.6167 | 0.67358 | 4.0987 | 0.93062 | 1.3662
| 2.5333 | 0.67715 | 4.0988 | 0.93285 | 1.3663
| 2.6 | 0.67715 | 4.099 | 0.93106 | 1.3663
| 2.5833 | 0.67358 | 4.0993 | 0.9315 | 1.3664
| 2.5 | 0.67715 | 4.0995 | 0.93373 | 1.3665
| 2.4833 | 0.67715 | 4.0997 | 0.93418 | 1.3666
| 2.5667 | 0.67715 | 4.0998 | 0.93194 | 1.3666
| 2.55 | 0.67358 | 4.1 | 0.93238 | 1.3667
| 2.5333 | 0.67358 | 4.1004 | 0.93283 | 1.3668
| 2.45 | 0.67358 | 4.1005 | 0.93506 | 1.3668
| 2.5167 | 0.67715 | 4.1009 | 0.93327 | 1.367
| 2.4333 | 0.67358 | 4.101 | 0.9355 | 1.367
| 2.5 | 0.67358 | 4.1013 | 0.93371 | 1.3671
| 2.4833 | 0.67715 | 4.1017 | 0.93415 | 1.3672
| 2.4667 | 0.67715 | 4.102 | 0.93459 | 1.3673
| 2.3833 | 0.67715 | 4.1021 | 0.93683 | 1.3674
| 2.45 | 0.67358 | 4.1024 | 0.93503 | 1.3675
| 2.3667 | 0.67358 | 4.1025 | 0.93727 | 1.3675
| 2.4333 | 0.67715 | 4.1027 | 0.93548 | 1.3676
| 2.35 | 0.67715 | 4.1029 | 0.93771 | 1.3676
| 2.4167 | 0.67715 | 4.1031 | 0.93592 | 1.3677
| 2.3333 | 0.67715 | 4.1032 | 0.93815 | 1.3677
| 2.3167 | 0.67715 | 4.1036 | 0.93859 | 1.3679
| 2.4 | 0.68072 | 4.1037 | 0.93636 | 1.3679
| 2.3833 | 0.67358 | 4.104 | 0.9368 | 1.368
| 2.3667 | 0.67358 | 4.1044 | 0.93724 | 1.3681
| 2.2667 | 0.67358 | 4.1046 | 0.93992 | 1.3682
| 2.35 | 0.67358 | 4.1047 | 0.93769 | 1.3682
| 2.25 | 0.67715 | 4.1051 | 0.94036 | 1.3684
| 2.3333 | 0.67715 | 4.1051 | 0.93813 | 1.3684
| 2.3167 | 0.67358 | 4.1054 | 0.93857 | 1.3685
| 2.2333 | 0.67358 | 4.1056 | 0.9408 | 1.3685
| 2.3 | 0.67715 | 4.1057 | 0.93901 | 1.3686
| 2.2167 | 0.67715 | 4.1059 | 0.94124 | 1.3686
| 2.2833 | 0.67001 | 4.1062 | 0.93945 | 1.3687
| 2.2667 | 0.67001 | 4.1067 | 0.9399 | 1.3689
| 2.1667 | 0.67358 | 4.107 | 0.94257 | 1.369
| 2.25 | 0.67358 | 4.107 | 0.94034 | 1.369
| 2.2333 | 0.67715 | 4.1074 | 0.94078 | 1.3691
| 2.15 | 0.67715 | 4.1074 | 0.94301 | 1.3691
| 2.1333 | 0.67715 | 4.1076 | 0.94345 | 1.3692
| 2.2167 | 0.67358 | 4.1078 | 0.94122 | 1.3693
| 2.2 | 0.67715 | 4.1081 | 0.94166 | 1.3694
| 2.1833 | 0.67358 | 4.1084 | 0.94211 | 1.3695
| 2.1 | 0.67358 | 4.1086 | 0.94434 | 1.3695
| 2.0833 | 0.67715 | 4.1089 | 0.94478 | 1.3696
| 2.1667 | 0.67358 | 4.109 | 0.94255 | 1.3697
| 2.15 | 0.67358 | 4.1093 | 0.94299 | 1.3698
| 2.1333 | 0.67358 | 4.1097 | 0.94343 | 1.3699
| 2.1167 | 0.67358 | 4.1099 | 0.94387 | 1.37
| 2.0333 | 0.67715 | 4.1101 | 0.9461 | 1.37
| 2.1 | 0.67358 | 4.1105 | 0.94432 | 1.3702
| 2.0833 | 0.67715 | 4.1108 | 0.94476 | 1.3703
| 2 | 0.67358 | 4.1109 | 0.94699 | 1.3703
| 1.9833 | 0.67358 | 4.1112 | 0.94743 | 1.3704
| 2.0667 | 0.67358 | 4.1113 | 0.9452 | 1.3704
| 1.9667 | 0.67358 | 4.1116 | 0.94787 | 1.3705
| 2.05 | 0.67715 | 4.1117 | 0.94564 | 1.3706
| 1.95 | 0.67715 | 4.1121 | 0.94831 | 1.3707
| 2.0333 | 0.67358 | 4.1121 | 0.94608 | 1.3707
| 1.9333 | 0.67715 | 4.1124 | 0.94875 | 1.3708
| 2.0167 | 0.67715 | 4.1125 | 0.94653 | 1.3708
| 2 | 0.67358 | 4.1127 | 0.94697 | 1.3709
| 1.9167 | 0.67715 | 4.1128 | 0.9492 | 1.3709
| 1.9833 | 0.67358 | 4.1132 | 0.94741 | 1.3711
| 1.9 | 0.67715 | 4.1132 | 0.94964 | 1.3711
| 1.9667 | 0.67715 | 4.1135 | 0.94785 | 1.3712
| 1.8833 | 0.67358 | 4.1137 | 0.95008 | 1.3712
| 1.8667 | 0.67358 | 4.114 | 0.95052 | 1.3713
| 1.95 | 0.67715 | 4.1141 | 0.94829 | 1.3714
| 1.85 | 0.67358 | 4.1144 | 0.95096 | 1.3715
| 1.9333 | 0.67715 | 4.1144 | 0.94873 | 1.3715
| 1.8333 | 0.67358 | 4.1147 | 0.9514 | 1.3716
| 1.9167 | 0.67715 | 4.1148 | 0.94918 | 1.3716
| 1.8167 | 0.67358 | 4.115 | 0.95185 | 1.3717
| 1.9 | 0.67358 | 4.1153 | 0.94962 | 1.3718
| 1.8 | 0.67358 | 4.1155 | 0.95229 | 1.3718
| 1.8833 | 0.67358 | 4.1157 | 0.95006 | 1.3719
| 1.8667 | 0.67358 | 4.1158 | 0.9505 | 1.3719
| 1.7833 | 0.67715 | 4.1159 | 0.95273 | 1.372
| 1.85 | 0.67358 | 4.1163 | 0.95094 | 1.3721
| 1.7667 | 0.67715 | 4.1164 | 0.95317 | 1.3721
| 1.8333 | 0.67358 | 4.1167 | 0.95139 | 1.3722
| 1.8167 | 0.67358 | 4.117 | 0.95183 | 1.3723
| 1.7333 | 0.67715 | 4.1171 | 0.95406 | 1.3724
| 1.8 | 0.67358 | 4.1174 | 0.95227 | 1.3725
| 1.7167 | 0.67358 | 4.1176 | 0.9545 | 1.3725
| 1.7 | 0.67358 | 4.1179 | 0.95494 | 1.3726
| 1.7833 | 0.67715 | 4.118 | 0.95271 | 1.3727
| 1.7667 | 0.67358 | 4.1183 | 0.95315 | 1.3728
| 1.6833 | 0.67715 | 4.1184 | 0.95538 | 1.3728
| 1.6667 | 0.68072 | 4.1187 | 0.95582 | 1.3729
| 1.75 | 0.67715 | 4.1188 | 0.9536 | 1.3729
| 1.65 | 0.67358 | 4.1191 | 0.95626 | 1.373
| 1.7333 | 0.67715 | 4.1192 | 0.95404 | 1.3731
| 1.7167 | 0.67358 | 4.1196 | 0.95448 | 1.3732
| 1.6333 | 0.67358 | 4.1196 | 0.95671 | 1.3732
| 1.7 | 0.67715 | 4.1197 | 0.95492 | 1.3732
| 1.6167 | 0.67358 | 4.12 | 0.95715 | 1.3733
| 1.6 | 0.67358 | 4.1202 | 0.95759 | 1.3734
| 1.6833 | 0.67358 | 4.1203 | 0.95536 | 1.3734
| 1.6667 | 0.67358 | 4.1207 | 0.95581 | 1.3736
| 1.5833 | 0.67358 | 4.1208 | 0.95803 | 1.3736
| 1.5667 | 0.67358 | 4.121 | 0.95847 | 1.3737
| 1.65 | 0.67715 | 4.1212 | 0.95625 | 1.3737
| 1.6333 | 0.67715 | 4.1214 | 0.95669 | 1.3738
| 1.6167 | 0.67715 | 4.122 | 0.95713 | 1.374
| 1.6 | 0.67358 | 4.1222 | 0.95757 | 1.3741
| 1.5833 | 0.67715 | 4.1227 | 0.95802 | 1.3742
| 1.5667 | 0.67715 | 4.1231 | 0.95846 | 1.3744
| 1.4833 | 0.67715 | 4.1231 | 0.96068 | 1.3744
| 1.55 | 0.67358 | 4.1235 | 0.9589 | 1.3745
| 1.4667 | 0.67358 | 4.1235 | 0.96112 | 1.3745
| 1.5333 | 0.67715 | 4.1238 | 0.95934 | 1.3746
| 1.45 | 0.68072 | 4.124 | 0.96157 | 1.3747
| 1.5167 | 0.67358 | 4.1243 | 0.95978 | 1.3748
| 1.4333 | 0.67001 | 4.1244 | 0.96201 | 1.3748
| 1.5 | 0.67715 | 4.1247 | 0.96023 | 1.3749
| 1.4167 | 0.67001 | 4.1248 | 0.96245 | 1.3749
| 1.4833 | 0.67358 | 4.125 | 0.96067 | 1.375
| 1.4 | 0.67715 | 4.1251 | 0.96289 | 1.375
| 1.3833 | 0.67358 | 4.1254 | 0.96333 | 1.3751
| 1.4667 | 0.67715 | 4.1255 | 0.96111 | 1.3752
| 1.45 | 0.67715 | 4.126 | 0.96155 | 1.3753
| 1.4333 | 0.67358 | 4.1261 | 0.96199 | 1.3754
| 1.35 | 0.67358 | 4.1264 | 0.96422 | 1.3755
| 1.4167 | 0.67715 | 4.1266 | 0.96243 | 1.3755
| 1.3333 | 0.67715 | 4.1267 | 0.96466 | 1.3756
| 1.4 | 0.67715 | 4.127 | 0.96288 | 1.3757
| 1.3167 | 0.67358 | 4.1271 | 0.9651 | 1.3757
| 1.3833 | 0.67358 | 4.1275 | 0.96332 | 1.3758
| 1.3667 | 0.67715 | 4.128 | 0.96376 | 1.376
| 1.35 | 0.67358 | 4.1283 | 0.9642 | 1.3761
| 1.2667 | 0.67358 | 4.1284 | 0.96643 | 1.3761
| 1.3333 | 0.67358 | 4.1287 | 0.96464 | 1.3762
| 1.25 | 0.67358 | 4.1288 | 0.96687 | 1.3763
| 1.2333 | 0.67715 | 4.1291 | 0.96731 | 1.3764
| 1.3167 | 0.67358 | 4.1292 | 0.96509 | 1.3764
| 1.2167 | 0.67358 | 4.1296 | 0.96775 | 1.3765
| 1.3 | 0.67358 | 4.1296 | 0.96553 | 1.3765
| 1.2833 | 0.67358 | 4.13 | 0.96597 | 1.3767
| 1.2667 | 0.67001 | 4.1303 | 0.96641 | 1.3768
| 1.1833 | 0.67715 | 4.1304 | 0.96863 | 1.3768
| 1.25 | 0.67358 | 4.1308 | 0.96685 | 1.3769
| 1.1667 | 0.67358 | 4.1308 | 0.96908 | 1.3769
| 1.2333 | 0.67358 | 4.1311 | 0.9673 | 1.377
| 1.15 | 0.67715 | 4.1312 | 0.96952 | 1.3771
| 1.2167 | 0.67358 | 4.1315 | 0.96774 | 1.3772
| 1.1333 | 0.67715 | 4.1317 | 0.96996 | 1.3772
| 1.2 | 0.67715 | 4.1321 | 0.96818 | 1.3774
| 1.1833 | 0.67715 | 4.1325 | 0.96862 | 1.3775
| 1.1 | 0.67358 | 4.1325 | 0.97084 | 1.3775
| 1.1667 | 0.67358 | 4.1329 | 0.96906 | 1.3776
| 1.15 | 0.67358 | 4.1333 | 0.96951 | 1.3778
| 1.05 | 0.67715 | 4.1336 | 0.97217 | 1.3779
| 1.1333 | 0.67358 | 4.1337 | 0.96995 | 1.3779
| 1.1167 | 0.67358 | 4.1341 | 0.97039 | 1.378
| 1.0333 | 0.67715 | 4.1343 | 0.97261 | 1.3781
| 1.1 | 0.67358 | 4.1346 | 0.97083 | 1.3782
| 1.0167 | 0.67358 | 4.1347 | 0.97305 | 1.3782
| 1.0833 | 0.67715 | 4.1349 | 0.97127 | 1.3783
| 1 | 0.67715 | 4.135 | 0.97349 | 1.3783
| 0.98335 | 0.67715 | 4.1353 | 0.97394 | 1.3784
| 1.0667 | 0.67715 | 4.1354 | 0.97172 | 1.3785
| 1.05 | 0.67358 | 4.1358 | 0.97216 | 1.3786
| 1.0333 | 0.67358 | 4.1363 | 0.9726 | 1.3788
| 1.0167 | 0.67358 | 4.1367 | 0.97304 | 1.3789
| 0.93335 | 0.67358 | 4.1368 | 0.97526 | 1.3789
| 1 | 0.67358 | 4.1371 | 0.97348 | 1.379
| 0.91668 | 0.67358 | 4.1372 | 0.9757 | 1.3791
| 0.98335 | 0.67358 | 4.1375 | 0.97393 | 1.3792
| 0.90001 | 0.67715 | 4.1376 | 0.97614 | 1.3792
| 0.96668 | 0.67715 | 4.1379 | 0.97437 | 1.3793
| 0.88335 | 0.67358 | 4.1381 | 0.97659 | 1.3794
| 0.95001 | 0.67001 | 4.1383 | 0.97481 | 1.3794
| 0.86668 | 0.67358 | 4.1386 | 0.97703 | 1.3795
| 0.93335 | 0.67715 | 4.1386 | 0.97525 | 1.3795
| 0.85001 | 0.67358 | 4.1389 | 0.97747 | 1.3796
| 0.91668 | 0.67715 | 4.1393 | 0.97569 | 1.3798
| 0.83335 | 0.67715 | 4.1394 | 0.97791 | 1.3798
| 0.90001 | 0.67358 | 4.1396 | 0.97614 | 1.3799
| 0.81668 | 0.67358 | 4.1399 | 0.97835 | 1.38
| 0.88335 | 0.67358 | 4.1402 | 0.97658 | 1.3801
| 0.80001 | 0.67715 | 4.1403 | 0.97879 | 1.3801
| 0.86668 | 0.67358 | 4.1406 | 0.97702 | 1.3802
| 0.85001 | 0.67358 | 4.141 | 0.97746 | 1.3803
| 0.76668 | 0.67358 | 4.1412 | 0.97968 | 1.3804
| 0.83335 | 0.67001 | 4.1415 | 0.9779 | 1.3805
| 0.75001 | 0.67715 | 4.1416 | 0.98012 | 1.3805
| 0.81668 | 0.67715 | 4.1419 | 0.97834 | 1.3806
| 0.73335 | 0.67358 | 4.1421 | 0.98056 | 1.3807
| 0.80001 | 0.67715 | 4.1423 | 0.97879 | 1.3808
| 0.71668 | 0.67358 | 4.1426 | 0.981 | 1.3809
| 0.78335 | 0.67358 | 4.1427 | 0.97923 | 1.3809
| 0.70001 | 0.67358 | 4.143 | 0.98145 | 1.381
| 0.76668 | 0.67715 | 4.1432 | 0.97967 | 1.3811
| 0.68335 | 0.67715 | 4.1435 | 0.98189 | 1.3812
| 0.75001 | 0.67358 | 4.1437 | 0.98011 | 1.3812
| 0.66668 | 0.67715 | 4.1438 | 0.98233 | 1.3813
| 0.73335 | 0.67715 | 4.1442 | 0.98055 | 1.3814
| 0.65001 | 0.67358 | 4.1444 | 0.98277 | 1.3815
| 0.71668 | 0.67715 | 4.1447 | 0.981 | 1.3816
| 0.63335 | 0.67358 | 4.1449 | 0.98321 | 1.3816
| 0.70001 | 0.67715 | 4.1451 | 0.98144 | 1.3817
| 0.61668 | 0.67358 | 4.1453 | 0.98365 | 1.3818
| 0.68335 | 0.67358 | 4.1455 | 0.98188 | 1.3818
| 0.60001 | 0.67358 | 4.1457 | 0.9841 | 1.3819
| 0.66668 | 0.67715 | 4.146 | 0.98232 | 1.382
| 0.58335 | 0.67358 | 4.1462 | 0.98454 | 1.3821
| 0.65001 | 0.67715 | 4.1464 | 0.98276 | 1.3821
| 0.56668 | 0.67715 | 4.1469 | 0.98498 | 1.3823
| 0.63335 | 0.67715 | 4.147 | 0.98321 | 1.3823
| 0.55001 | 0.67715 | 4.1472 | 0.98542 | 1.3824
| 0.61668 | 0.67358 | 4.1475 | 0.98365 | 1.3825
| 0.53335 | 0.67358 | 4.1477 | 0.98586 | 1.3826
| 0.60001 | 0.67358 | 4.1479 | 0.98409 | 1.3826
| 0.51668 | 0.67358 | 4.1482 | 0.9863 | 1.3827
| 0.58335 | 0.67001 | 4.1484 | 0.98453 | 1.3828
| 0.50001 | 0.67358 | 4.1488 | 0.98675 | 1.3829
| 0.56668 | 0.67715 | 4.1489 | 0.98497 | 1.383
| 0.48335 | 0.67358 | 4.1491 | 0.98719 | 1.383
| 0.55001 | 0.67358 | 4.1494 | 0.98542 | 1.3831
| 0.46668 | 0.67715 | 4.1497 | 0.98763 | 1.3832
| 0.53335 | 0.67715 | 4.1498 | 0.98586 | 1.3833
| 0.45001 | 0.67358 | 4.1502 | 0.98807 | 1.3834
| 0.51668 | 0.67715 | 4.1503 | 0.9863 | 1.3834
| 0.50001 | 0.67358 | 4.1508 | 0.98674 | 1.3836
| 0.41668 | 0.67358 | 4.1513 | 0.98896 | 1.3838
| 0.48335 | 0.67715 | 4.1514 | 0.98718 | 1.3838
| 0.40001 | 0.67001 | 4.1517 | 0.9894 | 1.3839
| 0.46668 | 0.67358 | 4.1519 | 0.98763 | 1.384
| 0.38335 | 0.67358 | 4.1524 | 0.98984 | 1.3841
| 0.45001 | 0.67358 | 4.1525 | 0.98807 | 1.3842
| 0.36668 | 0.67358 | 4.1528 | 0.99028 | 1.3843
| 0.43335 | 0.67358 | 4.1531 | 0.98851 | 1.3844
| 0.35001 | 0.67715 | 4.1534 | 0.99072 | 1.3845
| 0.41668 | 0.67715 | 4.1536 | 0.98895 | 1.3845
| 0.40001 | 0.67715 | 4.154 | 0.98939 | 1.3847
| 0.31668 | 0.68072 | 4.1545 | 0.99161 | 1.3848
| 0.38335 | 0.67358 | 4.1547 | 0.98984 | 1.3849
| 0.30001 | 0.68072 | 4.155 | 0.99205 | 1.385
| 0.36668 | 0.67358 | 4.1552 | 0.99028 | 1.3851
| 0.35001 | 0.67358 | 4.1558 | 0.99072 | 1.3853
| 0.33335 | 0.67715 | 4.1562 | 0.99116 | 1.3854
| 0.26668 | 0.67715 | 4.1564 | 0.99293 | 1.3855
| 0.31668 | 0.67358 | 4.157 | 0.9916 | 1.3857
| 0.30001 | 0.67715 | 4.1576 | 0.99204 | 1.3859
| 0.21668 | 0.67715 | 4.1582 | 0.99426 | 1.3861
| 0.28335 | 0.67715 | 4.1583 | 0.99249 | 1.3861
| 0.26668 | 0.67715 | 4.1589 | 0.99293 | 1.3863
| 0.25001 | 0.67358 | 4.1595 | 0.99337 | 1.3865
| 0.18335 | 0.67358 | 4.1597 | 0.99514 | 1.3866
| 0.23335 | 0.67358 | 4.1601 | 0.99381 | 1.3867
| 0.16668 | 0.67715 | 4.1603 | 0.99558 | 1.3868
| 0.21668 | 0.67358 | 4.1608 | 0.99425 | 1.3869
| 0.15001 | 0.67715 | 4.1613 | 0.99602 | 1.3871
| 0.20001 | 0.67358 | 4.1614 | 0.9947 | 1.3871
| 0.13335 | 0.67715 | 4.162 | 0.99647 | 1.3873
| 0.18335 | 0.67358 | 4.1622 | 0.99514 | 1.3874
| 0.11668 | 0.67358 | 4.1628 | 0.99691 | 1.3876
| 0.16668 | 0.67358 | 4.163 | 0.99558 | 1.3877
| 0.15001 | 0.67358 | 4.1638 | 0.99602 | 1.3879
| 0.10001 | 0.67358 | 4.1639 | 0.99735 | 1.388
| 0.083347 | 0.67358 | 4.1647 | 0.99779 | 1.3882
| 0.13335 | 0.67358 | 4.1648 | 0.99646 | 1.3883
| 0.11668 | 0.67001 | 4.1656 | 0.99691 | 1.3885
| 0.066681 | 0.67715 | 4.1657 | 0.99823 | 1.3886
| 0.10001 | 0.67715 | 4.1665 | 0.99735 | 1.3888
| 0.050014 | 0.67715 | 4.1669 | 0.99867 | 1.389
| 0.083347 | 0.67358 | 4.1677 | 0.99779 | 1.3892
| 0.033347 | 0.67358 | 4.1683 | 0.99912 | 1.3894
| 0.066681 | 0.67715 | 4.1688 | 0.99823 | 1.3896
| 0.050014 | 0.67358 | 4.17 | 0.99867 | 1.39
| 0.016681 | 0.67715 | 4.1702 | 0.99956 | 1.3901
| 0.033347 | 0.67358 | 4.1716 | 0.99912 | 1.3905
| 0.016681 | 0.67715 | 4.1737 | 0.99956 | 1.3912
| 1.3889e-05 | 0.67358 | 4.1774 | 1 | 1.3925
| 1.3889e-05 | 0.67358 | 4.1802 | 0.99999 | 1.3934
| 0 | 0.67358 | 4.2 | 1 | 1.4

## Original (hidden)
total capacity | 17.5 | mAh | SOC | voltage
shelf time | 120 | days | 0 | 3
shelf current | 50 | nA | 1 | 3.27
self discharge rate | 10% | percent per year | 2 | 3.39
shelf energy | 0.144 | mAh | 3 | 3.455
shelf capacity | 4.111% | SOC | 5 | 3.558
HR current | 2.06 | mA | 10 | 3.658
HR time | 600 | s | 15 | 3.676
HR energy | 0.34333 | mAh | 20 | 3.708
HR capacity | 1.962% | SOC | 25 | 3.728
required for HR measure | 40% | SOC | 30 | 3.743
fitting time | 30 | min | 35 | 3.757
fitting current | 1.06 | mA | 40 | 3.767
fitting energy | 0.53 | mAh | 45 | 3.782
fitting capacity | 3.029% | SOC | 50 | 3.797
capacity at EOL | 49.1% | SOC | 55 | 3.817
capacity rounded up at EOL | 50% | SOC | 60 | 3.845
voltage at EOL | 3.797 | V | 65 | 3.876
capacity at beginning of fitting | 45.89% |  | 70 | 3.91
capacity after fitting | 42.86% |  | 75 | 3.95
|  |  | 80 | 3.988
|  |  | 85 | 4.032
|  |  | 90 | 4.075
|  |  | 95 | 4.12
|  |  | 100 | 4.18